NIH Image: Use In Fluorescence And Confocal Microscopy

Version 2.0

October, 1996

Prepared by

Harvey J. Karten

Dept. of Neurosciences

University of California, San Diego

La Jolla, CA 92093-0608

EMail: hjkarten@ucsd.edu

Table of Contents

NIH Image: Use In Fluorescence And Confocal Microscopy	1

EMail: hjkarten@ucsd.edu	1

Table of Contents	2

Acknowledgments	5

I.	INTRODUCTORY COMMENTS	6

1.	NIH Image On Macintosh And PC/Windows 95	6

2.	Contents Of This Manual	6

3.	Some Hardware Considerations	7

A)	Operating System	8

B)	Graphic RAM and Monitors	8

C)	RAM vs. Virtual Memory	8

D)	Disks And Storage Media	8

E)	Extensions	9

4.	Loading NIH Image And The Confocal Macros	9

A)	Macros	9

5.	Other Software Used In Conjunction With NIH Image	9

A)	Adobe Photoshop 3.0.5 And Canvas 5.0	10

B)	FileMaker Pro 3.0	10

6.	Image Database Software	10

7.	Cataloging Software: Where Are The Files When You Need Them?	10

II.	EVALUATING FLUORESCENT IMAGES WITH NIH IMAGE PRIOR TO CONFOCAL IMAGING	11

1.	Evaluation of Fluorescent Images Prior to Confocal Microscopy	11

2.	Fluorescent Excitation/Emission: A Moving Target	11

3.	Selection Of Fluorophores	12

4.	Using NIH Image To Visualize Cy5	13

5.	On-Chip Integration	13

6.	Practical Guidelines For Implementing On-Chip Integration	14

7.	Macros For Shutter Control	14

8.	On-Chip Integration, Multiple Labeled Sections, And RGB Images	15

III.	OPENING CONFOCAL IMAGES IN NIH IMAGE	16

1.	Transferring Files From A DOS, DOS/Windows Or OS/2 Based Computer To A Macintosh	16

A)	Sneaker Net	16

B)	Local Area Network (LAN)	17

C)	FTP Via Internet	17

D)	ZIP Drives	18

1.	A Few Words Of Caution About Using Zip Disks	18

E)	Magneto-Opticals Disks For Archival Storage	18

2.	Opening BioRAD Files In NIH Image	19

A)	BioRAD Split Screen Images	20

B)	Merging Split Screen Z-Series	21

3.	Leica Files	21

4.	Zeiss Files	22

5.	Molecular Dynamics/Sarastro	22

6.	Noran Confocal Files	23

IV.	BASIC IMAGE MANIPULATIONS	23

1.	Evaluating The Quality Of Your Original CLSM Image	23

A)	Making Sure That Your Original CLSM Image Uses The Full Range Of 8-Bit Values (0-255)	24

B)	Avoiding High Contrast Images	24

C)	Avoiding Noisy Images	24

2.	Editing Image	24

A)	Cropping Images, Erasing, Superimposing Text, Scale Bars, Rotating And Scaling Images	24

1.	Notes On Scaling	25

2.	Caution On Scale And Rotation Of Images	25

3.	Using LUTs	25

A)	Modifying Brightness And Contrast	25

B)	Linear And Non-Linear LUTs, Including Custom LUTs.	25

C)	Enhance Contrast Operator In NIH Image	26

D)	Thresholding And Density Slicing	26

E)	Pseudocoloring Images.	26

F)	Exporting To Adobe Photoshop	26

4.	Enhancing/Filtering Image	27

5.	Quantitative Measurements	27

6.	NIH Image Macro Language	27

V.	ADVANCED TOPICS	27

1.	Merging Pairs Of Double Labeled Sections	27

A)	NIH Image and Adobe Photoshop	28

B)	Double Labeled Sections: Building A Stack (Best Done Using A Macro)	29

1.	Color Merge In NIH Image	29

2.	Merging Files With Adobe Photoshop	30

3.	Preferred Method	31

C)	Compare Results Of NIH Image 8-Bit Merge With Adobe Photoshop 24-Bit Merge	31

D)	Adjusting Color Contrast On Sections In A Stack (See Macros)	31

E)	Merging A Double Labeled Pair Of Z-Series Using A Macro.	31

2.	Projection Of A Z-Series And 3D Rotations	32

A)	Stepping Through A Z-Series Using Stacks.	32

B)	Animating A Z-Series	32

C)	Projecting A Z-Series Onto A Single Plane	32

1.	Selection Of An Area For Z-Projection	32

2.	Optimizing Settings For Project Function Of Z-Series	34

D)	Reslicing The Z-Series Along Alternate Planes: (X-Z, Y-Z And Theta-Z)	35

E)	Reslicing To Make A Z-Series In An Alternate Plane	35

1.	Rapid Dynamic 3D Reslicing	36

F)	Generating A 3D-Series From A Z-Series	36

G)	Animating A 3D-Series (Producing "Apparent Rotation")	37

H)	Make A Stereo Pair Or Series Of Stereos?	37

1.	Stereo Series In Black And White	38

2.	Stereo Pair Of Single Labeled Section In Color	38

3.	Color Stereo Images Of A Double Labeled Section	39

I)	Exporting Stacks To QuickTime Movies And VCR Recordings Of Stack	39

1. Saving Stacks In QuickTime Format	39

VI.	PRINTING ON VIDEO PRINTER, DYE SUBLIMATION PRINTER, SLIDE MAKER AND VCR	40

1.	Video Printer	40

2.	Dye Sublimation Printer	40

3.	Slide Maker	40

4.	VCR	41

VII.	USE OF NIH IMAGE FOR IMAGE COLLECTION AND INSTRUMENT CONTROL	41

VIII.	APPENDIX	41

1.	Macros	41

IX.	ADDENDA TO NIH IMAGE CONFOCAL IMAGING MANUAL	41

Addenda to Manual March 1995	41

Addenda June, 1995	41

Addenda September 1996	42

List of Macros provided in Confocal Macros	42

Define Procedures for Stacks	42

Some Universal Operations	42

Operations on Confocal Files: Import, Split, Merge	42

General Stack Operations	42

Shutter Controls and On-Chip Integration	43

Stack Modifications, Z-Projections, Stereo Pairs	43

LUT Manipulations	44

Hot Keys	44

Macros	45

Acknowledgments

With many thanks to Wayne Rasband for his invaluable contributions to scientists of all disciplines, and his advancement of image processing.

The work that led to this manual was supported by grants to H. J. Karten from NIMH, NINDS and NEI. This manual  was prepared with the direct support of NIMH and the Human Brain Project and the Brain Database Project. Under the stress of attempts to obtain grant support, we often neglect to sufficiently thank the many dedicated scientists who work so hard to establish those programs that provide the necessary funds. This manual is dedicated to those many loyal supporters of all of us at NIH, NSF, DOD, NASA, and elsewhere.

I have adopted and adapted a number of macros contributed by various NIH Image users. The authorship of the original macros was often difficult to determine. I am grateful to the unacknowledged authors, and hope that they are not offended if not specifically credited. My special thanks to Rusty Gage and Larry Goldstein for providing me with unlimited access to their confocal microscope facilities. Without their generosity, this manual would not have been possible.

Please let me know if this manual is helpful. Please inform me of any of errors in this manual. If you have any suggestions to improve this manual, additional macros that are helpful for processing confocal images, or other strategies for processing CLSM images, please let me know, and I will try to incorporate these changes in future versions.

Happy confocaling,

Harvey J. Karten, M.D.

hjkarten@ucsd.edu

I. INTRODUCTORY COMMENTS

NIH Image is a useful tool for evaluating fluorescent material prior to examining it on the confocal microscope and for post-processing of confocal laser scanning microscope (CLSM) images. The experienced user of NIH Image may find many of these operations obvious. However, many users of CLSMs, previously unfamiliar with NIH Image will find it a useful tool for both pre- and post-acquisition processing of images. I hope that this manual provides a relatively easy introduction to the use of NIH Image for CLSM users.

Performing post-acquisition processing of images on the Macintosh will free the confocal instrument for image collection. Though not as extensive as VoxelView, Analyze, VolVis, SYNU or VoxBlast, particularly in applications requiring advanced volume rendering and Voxel based calculations, the use of NIH Image and Adobe Photoshop 3.0 (Adobe Systems Inc., Mountain View, CA) on the PowerPC will provide most of the functions needed for manipulating confocal images, and at a much lower price.

This manual is intended for use with NIH Image, version 1.60 or later. Several of the procedures described in the manual will not operate correctly on earlier versions. You can obtain a copy of the latest version of NIH Image from the FTP site: zippy.nimh.nih.gov (login: anonymous; Password: <your EMail address>; cd/pub/image) or the associated Website (http://rsb.info.nih.gov/nih-image/).

1.	NIH Image on Macintosh and PC/Windows 95

This manual describes the use of NIH Image on a Macintosh. In principle, it should be fully applicable to a version of NIH Image being prepared for PC/Windows 95.

Tod Weinberg of Scion Corporation (Frederick, MD) is sponsoring the transfer of NIH Image to PC/Windows95. This is a direct "port" of NIH Image--i.e., it should have the exact look and feel of NIH Image for the Macintosh, with the exception of those features of the operating system that reflect the individual platforms. At this time (October, 1996), ImagePC (NIH Image for PC) is still in alpha testing, and some operations are not yet functional, but it is available for downloading by interested individuals. It is available from the same location that provides NIH Image for the Macintosh, both from the FTP site and the Website. The "port" of the program to the PC seems very promising, and the final version will hopefully permit users to perform the same operations on the PC and Macintosh with equal facility.

2.	Contents Of This Manual

A)	How to use NIH Image to evaluate the quality of single or multiple labeled fluorescent histological sections in preparation for confocal microscopy.

B)	How to transfer files from the BioRAD, Leica or Zeiss CLSM to NIH Image, both individual images and Z-series.

C)	How to use NIH Image for image processing and analysis: NIH Image provides a wide range of functions for image analysis and processing, including changing contrast and brightness values, pseudocoloring images, rotating, cropping, scaling, various filtering operations, measuring density, density

slicing, measuring length and area, profile of a line, and counting particles. These are useful tools for modifying and analyzing confocal images. These functions fall into the more general category of image processing, and are discussed at greater length in the NIH Image manual, About NIH Image.

D)	How to use multiple windows, Z-series and Stacks: The Macintosh allows you to open many windows, each containing a different confocal image. This facilitates comparison of different images. The various Stacks functions are amongst the most powerful features of NIH Image. NIH Image Stacks function allows you to rapidly step back-and-forth through a Z-series of sections. You can crop this Stack to select specific features of interest. You can also use the Stacks function to rapidly and alternately compare different images.

E)	How to "Project" all the plates (or only a selected number of plates) of a Z-series onto a single plane.

F)	How to generate 3D images and 3D rotations: NIH Image allows you to generate a 3D-series of images from the Z-series. You can display them in a montage window, fabricate stereo pairs or animate them to give the appearance of rotating a 3D object in space.

G)	How to merge double and triple labeled sections, producing a three color images (RGB) with NIH Image.

H)	How NIH Image can generate a double/triple labeled 3D-series for stereo pairs and rotations.

3.	Some Hardware Considerations

Macintosh computers have proven particularly suitable for graphics applications. Their 32-bit memory model permits use of large quantities of RAM for rapid processing of image files. The graphics hardware and operating system have proven very suitable for image processing. All the procedures described in this manual can be achieved on a "low end" Macintosh, such as the Mac IIci, IIfx, to midlevel machines such as the 68040 Quadra series and, of course, will perform extremely well on the PowerPC series of machines. Most of the routines described in this manual were developed using a Quadra 950. Separate versions of NIH Image, optimized for the 680X0 series, as well as a single "fat" binary version that is optimized for both a 680X0 and a PowerPC, are available. Both versions are now available from the FTP site. Much less expensive than many dedicated graphics workstations, such as a Silicon Graphics Indy^2, the cost of a fully equipped Macintosh will vary in price based on some of the variables listed below. NIH Image will perform adequately on any current Macintosh.

Major differences in performance will be noticed when generating a rotation series, resectioning a Z-series, or applying various filters. It is not necessary to have a 68040 with a floating point processor (FPU). However, for improved speed performance, a PowerPC 7100, 8500, or the 9500, will prove most attractive. The 6100 does not have room for a frame grabber or a video output card to video printers or VHS tape. If you have the funds, the 9500/200 MHz unit is obviously preferable. The 8500 and 9500 have fast SCSI-2 ports for rapid data file transfers between external devices, built-in Ethernet for communication, and can be equipped with a high speed graphics display card. NIH Image does not support a multiprocessor Macintosh.

As we accumulate more experience using NIH Image on a PC/Windows 95 platform, future editions of this manual will include comments pertinent to that platform. Based on preliminary considerations, we suggest the use of a fast Pentium (e.g., Pentium Pro 200), with large quantities of RAM, fast 24-bit graphics card, a large hard disk and a 20-inch RGB monitor.

A)	Operating System

Use System 7.1 or higher. System 7.5.5 is quite stable, and we now use it routinely.

B)	Graphic RAM and Monitors

NIH Image is an 8-bit program. However, beginning with version 1.56 NIH Image will run even if the monitor is in 24-bit mode. This permits rapid switching between Adobe Photoshop 3.0 (24-bit program) and NIH Image. In order to obtain a full 24-bit image on a 20 inch monitor (at 1024x768 or higher), the graphics display card on an PowerPC should be packed with 4 MB of VRAM to permit 24-bit color on a 20-inch monitor. The Apple Trinitron 20-inch multisynch monitor provides an excellent image, and also is the lowest priced 20-inch monitor for the size and quality of the image. However, all the operations outlined in this manual will work on a 14-inch monitor with 8-bits. The 7100 and 6100 only permit 2 MB of VRAM, thus you will be limited to using a 17-inch monitor with 24-bit color.

C)	RAM Vs. Virtual Memory

You will need large quantities of RAM for optimal performance. I suggest a minimum of 32 MB of RAM, and more if you can afford it. Avoid Virtual Memory if at all possible. (However, Apple does recommend that you assign 1 MB of Virtual Memory when using System 7.5.x on a PowerPC.) Graphics files are large, and confocal Z-series are often huge. If you also have a motorized stage and make extended XY planes of Z-series, your files are likely to be 30-50 MB or larger and you will have to reconcile yourself to buying a minimum of 80-256 MB of RAM if you want to move faster than a glacier. A minimal rule of thumb is that you should have 2.5-3 times more RAM than the size of your largest files. You should also learn how to allocate this memory to NIH Image using Get Info. Otherwise the additional memory will be of no benefit to you. I suggest that you allocate a minimum of 24 MB of RAM to NIH Image, and another 24 MB to Adobe Photoshop. Play around with the program. Many people are using RAM Doubler to compensate for limited amounts of RAM. I have no experience with this program and have heard mixed reports as to it value in working with large graphics files.

D)	Disks And Storage Media

You will need storage media with a large capacity. But I assume that if you are working with confocal images, you have already had to deal with this problem. A hard disk of at least 500 MB is required. For greater flexibility, a system disk of 1 GB and a magneto- optical disk of 500 MB or larger for your data files is suggested. You may find it helpful to transfer your larger files from the slower magneto-optical to your system hard disk when working on them. Then move them back to the magneto-optical for long term storage. Confocal Z-series may readily exceed 25 MB. You should have enough storage space on your hard disk to allow you to save both the original file and any

modifications you may make during a work session. We often find that we use more than 75 MB per Z-series in a single session.

E)	Extensions

A few words about those clever System Extensions on the Macintosh--the Extensions that make tea while you compute, or respond to voice commands (such as Apple’s PlainTalk). These extensions not only chew up space, but worse, they convert your PowerPC to molasses while continually polling for voice inputs, losing CPU cycles, etc. Turn them all off. Stay with the basics. When you are doing image processing, turn off "File Sharing". If someone else on the network decides to just glance at your directory, your machine will be distracted and slow things down. You will still be able to get back onto the network with the click of key, but not be impeded by curious onlookers. Get used to dedicated computing, as in the days of yore, if you want to obtain your results as speedily as possible.

4.	Loading NIH Image And The Confocal Macros

This manual is written with the assumption that the reader is familiar with the Macintosh Operating System 7.1 or higher.

Make sure that you have allocated sufficient memory to NIH Image to perform many of the operations outlined. In addition to allocating memory to the program as described above, using the File menu item, Get Info, you must also allocate sufficient memory to the NIH Image "Undo & Clipboard Buffer Size", shown under the Options + Preferences menu item. I suggest a minimum of 1400 KB. You must then Record Preferences under the File menu. Close the program and Restart the program.

A)	Macros

The accompanying file, "Confocal Macros" should be copied into the Folder (Directory) containing the other NIH Image macros. Select the Specials menu, and the Load Macros item. Select the "Confocal Macros" from the resulting dialog box.

The “Confocal Macros” file contains a series of macros that are particularly useful for processing double labeled pairs of sections, Z-series, and generating stereo pairs. The reader should also familiarize themselves with the use of Stacks in NIH Image, as described in About NIH Image.

A copy of the macros file is included in the binhexed file "/pub/image/documents/confocals.hqx" provided via FTP. However, for those users who have obtained this file directly from the Website, the confocal macros file is not readily available. In order to provide access to this file, "Confocal Macros" has been appended to the end of this manual. It should be copied to a "Simple Text" file, and saved as "Confocal Macros."

5.	Other Software Used In Conjunction With NIH Image

NIH Image provides a variety of powerful operations not readily available in other software, regardless of the price. Some functions, however, are best performed using other software packages. This is particularly true for those procedures that benefit from true 24-bit operations, such as adjusting individual color planes, color filtration and printing 24-bit images to dye sublimation printers.

A)	Adobe Photoshop 3.0.5 and Canvas 5.0

Most confocal microscopy labs currently use Adobe Photoshop 3.0 for adjusting color saturation, brightness/contrast, cropping, assembling composite illustrations of multiple images, labeling the images and printing RGB confocal images. A new version of Canvas 5.0 (Deneba Corporation, Miami, FL) provides all these important functions of Adobe Photoshop needed by confocal microscopists for about one-third of the cost. Canvas 5.0 handles text labels and layers much better than Adobe Photoshop, but the cropping tool takes a bit of getting used to.

B)	FileMaker Pro 3.0

In order to maintain a record of the content of each image or series of images as you collect them on the confocal scope, I have prepared a FileMaker Pro template. The Template is also posted on the FTP server for NIH Image within "confocals.hqx". When you download and decompress "confocals.hqx" the template will be decompressed as "Confocal_FileMaker_Template". FileMaker Pro is a simple and inexpensive database program. It is available for both Macintosh and DOS/Windows, and files are interchangeable between the two versions.

The careful design of a database will encourage you to store important information including the material and parameters used for the collection of the original image, storage location of the file, and modifications to the file.

6.	Image Database Software

FileMaker Pro 3.0 can save images as well as text. However, dedicated image databases will automatically scan a disk and generate thumbnail images, store location, and can directly reopen the original image file. Many users may find this preferable. There are several useful image databases, including Multi-Ad Search 3.1 (Multi-Ad Services, Peoria, IL), Cumulus 2.5, Kudo and Kodak Shoebox (Adobe Fetch is no longer available). Of the various databases that I have tested, Multi-Ad Search 3.1 and Cumulus 2.5 are the most useful. Cumulus is particularly advantageous as it supports Apple ScriptMaker, and can automatically transfer thumbnail images to a FileMaker Pro database. A new low cost single-user version of Cumulus (3.0) is scheduled for release in the Fall of 1996.

7.	Cataloging Software: Where Are The Files When You Need Them?

By this time, you will be overloaded with files, different disks, floppies, Zip, Jaz, magneto-opticals, Syquests and everything but paper tape. You will have various versions of the same file in many different locations. How will you ever find a file that you need for a publication?

Iomega Corp. (Roy, UT) provides a utility called "FindIt" with their Jaz drives that makes a compiled directory of the contents of various hard disks and removables. It works very well with magneto-opticals disks, Zip and Jaz drives, and floppies. However, it will only work with drives that have Macintosh formatting. It does not even recognize the presence of a PC-formatted Zip or PC-formatted magneto-optical disk (see below).

II.	EVALUATING FLUORESCENT IMAGES WITH NIH IMAGE PRIOR TO CONFOCAL IMAGING

1.	Evaluation Of Fluorescent Images Prior To Confocal Microscopy

The best way to obtain excellent final confocal images is to start with a good specimen. All the digital magic in the world won't make a good picture if you start with a bad specimen.

Careful evaluation of your fluorescent labeled specimen on a high quality fluorescent microscope will save you a lot of time and effort on the confocal microscope. Your best confocal images will be obtained from sections of highest quality. Such sections are easily identified on a standard fluorescent microscope. The qualities that should be evaluated include:

A)	General tissue quality, including fixation, lack of tears or folding in tissue

B)	High signal to noise ratio. The labeled processes should be bright and readily distinguished from the background. High background fluorescence poses a difficult problem that cannot be easily overcome by even the best confocal microscope. Poor signal to noise ratio will prompt you to modify your final image by increasing the contrast to excess.

C)	Good separation of fluorophores. Very bright intensity of fluorescence of tetramethyl rhodamine isothiocyanate (TRITC) or indocarbocyanine (Cy3) will produce substantial breakthrough of the image into the range of the fluorescein isothiocyanate (FITC) image. While your naked eye will note this to be somewhat red in color, the PMT in the CLSM is colorblind and will detect this and not distinguish it from the FITC component of the image. While additional filters may help, there are several strategies that will help you deal with this problem.

1.	Do not use TRITC with FITC. Try rhodamine sulfonyl chloride (LRSC) or indodicarbocyanine (Cy5) in conjunction with an Argon-Krypton laser.

2.	Examine a single stained section with each fluorophore. If you see breakthrough from LRSC into the FITC filter set, you know you have trouble.

3.	Swap fluorophores. If the FITC label was attached to an antigen in low concentration, and the LRSC to that in high concentration, this will increase the prospects of breakthrough.

D)	If you are examining the distribution of two or more substances in the section, your strategy will depend upon whether the two antigens are in the same or in different loci in the tissue. If one antigen is in the nucleus, and the other in the cytoplasm, then you obviously don't have to worry about minor levels of breakthrough.

2.	Fluorescent Excitation/Emission: A Moving Target

Fluorescent excitation/emission is a degradative process. The brighter the excitation (within limits), the brighter the resulting emitted image. In order to examine the image for adequate evaluation of its quality and content, the user will expose the tissue for

increasing lengths of time. Both excitation brightness and duration of exposure, while required for evaluation of the tissue, also degrade the quality of the fluorescence. At saturating levels, the half life of FITC and TRITC is only about 1-2 seconds. When dealing with in vivo or in vitro specimens, the problem is even more notable, as the prolonged exposure to intense excitation light degrades the cells, and hastens cell death. (This has been used to good benefit in the selective killing of identified cells in nervous tissue). Achieving good dark adaptation by the observer, use of efficient microscopes, supercooling the specimen, pulsing the light source and various reagents (often toxic to living cells) have all been used to diminish the fading, but the inevitable lose of fluorescence cannot be fully avoided.

As a result of these limitations, you may find that you are reluctant to examine the tissue at length before going to the confocal microscope to capture your "perfect” confocal images. This results in many bad confocal images, often consequent to the fact that the original tissue was not properly evaluated before using the CLSM. All the tricks in the world of post-acquisition processing of confocal images won't improve a lousy specimen. An expensive CLSM won't take a good picture of a bad specimen.

Far too often, investigators would be better served learning how to optimize their use of a fluorescent microscope more thoroughly before spending time hacking images from a confocal microscope.

3.	Selection Of Fluorophores

This is a large and complex topic. It is dealt with extensively in the excellent chapter by Brelje, Wessendorf and Sorenson (1993) in the monograph Methods in Cell Biology: Cell Biological Applications of Confocal Microscopy, edited by Brian Matsumoto.

If working with a single fluorophore, most users seem to prefer either FITC or Cy3. FITC has long been used, and the emission wavelength corresponds to the range of peak sensitivity of the human eye. Cy3, with an emission wavelength in the short red range, has become increasingly popular. Cy3 is intensely fluorescent and bleeds both high and low--use it only in single fluorophore configuration, unless you have previously determined that the two compounds of interest definitely do not co-localize in the same structure. The great advantage of so bright a fluorophore is that you can use a smaller aperture, with a favorable signal to noise ratio, lower laser power and shorter exposure times.

My own preferences are:

A)	Single fluorophore: Cy3

B)	Double label fluorophores: FITC and LRSC or FITC and Cy5

C)	Triple label fluorophores: FITC, LRSC and Cy5

D)	Quadruple label: aminomethylcoumarin acetate (AMCA), FITC, LRSC and Cy5 (AMCA requires a laser that excites in the range of UV to short blue. While expensive, microscopes with lasers in this range are increasingly available.)

All these fluorophores can be obtained from Jackson Immunoresearch Laboratories, Incorporated (West Grove, PA).

4.	Using NIH Image To Visualize Cy5

Many people have difficulty with Cy5 for the very reason that makes it so useful--it's emission is widely separated from that of FITC, and thus is barely visible to the naked eye. It is only visible if you are completely dark adapted, have a very efficient microscope and the intensity of fluorescence is robust. Even then, you will not be able to see much of the fine detail.

However, as users of NIH Image, you have one of the best tools available for imaging Cy5, if you have a video camera that supports on-chip integration (see following section) and a Scion LG-3 frame grabber. CCD video cameras are quite sensitive in the red to infra-red range. Wayne Rasband has provided a useful macro that allows you to vary the on-chip integration time and adjust the duration of integration while viewing the image on your computer monitor. For even greater control of signal fading, you should also use a Uniblitz shutter between the fluorescent light source and the microscope (see "Macros For Shutter Control" below).

5.	On-Chip Integration

The intensity of fluorescent images is relatively low compared to the sensitivity of most CCD cameras. Video cameras sample available light, then send a complete image to the monitor every 1/30 of a second (33.3 msec or 30 frames per second). At that moment, the previous image is erased and the camera again starts to accumulate a charge proportional to the intensity of the incident light.

As the amount of light declines, the camera produces a more grainy, noisy image. With progressively lower light levels, the camera fails to detect a sufficient number of incident photons to provide a useful image. Averaging frames will not help, nor will software integration of multiple weak images.

In photography with film, the user has a number of options:

A)	Increase the speed of the film

B)	Increase the aperture or light collecting ability of the optics

C)	Increase the exposure time--i.e., decrease the shutter speed

The analogous situation pertains to improving image capture with video cameras.

A)	Increase the speed of the film

If you need to observe rapidly changing events, the only strategy you have available is to increase the sensitivity of your detector by using an intensified video camera. This is quite expensive, and the quality of the image is often degraded by the intensifier.

B)	Increase the aperture or light collecting ability of the optics

There have been great improvements in the quality of fluorophores, light sources, lenses and filters. However, many images are still too faint to be adequately captured with a standard video camera.

C)	Increase the exposure time--i.e., decrease the shutter speed

If your specimen is stable, then you can opt for longer exposure times.

On-chip integration acts by increasing the exposure time, thus increasing the number of photons captured. Rather than sending the image to the monitor every 33 msec, photons are allowed to accumulate on the sensor, and the image sent to the monitor when it is deemed adequately saturated. This technology has been available for many decades, but the cost often seemed beyond the budget of most labs.

NIH Image, in conjunction with some inexpensive hardware, now permits this technology to be widely used.

In order for this to work properly, you need four components:

A)	A video camera capable of on-chip integration.

B)	A source for a properly timed TTL pulse.

C)	A frame grabber (video digitizer) that grabs the image at exactly the correct moment that the camera is sending the integrated image.

D)	A computer and software that can control this process.

6.	Practical Guidelines For Implementing On-Chip Integration

A)	Choice of Camera: Several video cameras now provide built-in circuitry for on-chip integration at no additional cost. These include the two most widely used video cameras in labs, the Cohu 4915 (Cohu Inc., San Diego, CA) and the Dage-MTI 72 (Dage-MTI, Inc, Michigan City, IN). Images are allowed to accumulate on the camera sensor until a suitable level of exposure is achieved. A TTL pulse is then sent to the camera indicating that the computer now expects to receive the accumulated (integrated) image.

B)	and C) TTL Pulse Source and Synchronized Frame Grabber: The Scion LG-3 card has made this technology available at low cost. Under normal operating conditions, the LG-3 continually samples the input from the video camera and displays the results to the monitor. However, under appropriate software control, the LG-3 card provides an output of a TTL pulse that is synchronized to the video digitizer, and captures the next image and holds it in the display buffer.

D)	The software for controlling this whole process is provided in NIH Image. Wayne Rasband has provided a macro to control the duration of the exposure time for on-chip integration.

7.	Macros For Shutter Control

I strongly recommend the use of Uniblitz shutter between the excitation light source and the specimen. The Uniblitz shutter is manufactured by Vincent Associates in Rochester, N.Y. Most microscope manufacturers can provide an adapter to mount the shutter between the excitation source and the microscope. The shutter is controlled by a Uniblitz controller box connected to the computer.

The advantage of using the shutter control is that it will limit the exposure of your tissue to the fading effects of the excitation by restricting exposure times to only that period required for actually collecting the images. Without this, I often forget to manually close the shutter, only to discover that I have burned out my best specimens.

Chi-Bin Chien, formerly of UCSD, has written a series of macros to control the shutter. I have combined the shutter control macros with Wayne's macro for on-chip integration. These macros are available as part of the collection of "Confocal Macros" provided at the NIH Image FTP site.

The following description pertains to the use of a fluorescent specimen. I suggest that you practice the use of on-chip integration with a transmitted light specimen, with the illumination set to a very low level. Once you have mastered the procedure for on-chip integration, you can then experiment with a fluorescent specimen.

A)	A camera window should be selected. The shutter can be closed. Turn off the AGC (automatic gain control) on the CCU (camera control unit ). Set the manual gain and manual black level fully counterclockwise.

B)	Select the macro for on-chip integration. If you do not have a Uniblitz shutter, you should open the manual shutter to allow excitation of your specimen.

C)	If the image is too dim, move the mouse towards the top of the camera window.

D)	If the image is too bright, move the mouse towards the bottom of the camera window.

E)	The information window in the lower left corner of the screen will report on the number of frames that were integrated for the latest image. You can modify the macro to provide a direct readout of the duration of integration in seconds.

F)	Observe the histogram for best spread of brightness values.

G)	You may find it helpful to make minor adjustments in the quality of the image by changing the gain and black level settings on the CCU.

H)	When you are satisfied with the quality of the image, terminate data collection by moving the mouse off the left side of the camera window and hold down the mouse button. The last image will continue to be displayed on the monitor. The shutter will then close, protecting the specimen from further fluorescent excitation.

This procedure requires practice to obtain images with a broad range of values.

As on-chip integration time increases, you will note a marked delay between the time you press down the mouse button and the appearance of a change in the image.

8.	On-Chip Integration, Multiple Labeled Sections, And RGB Images

NIH Image permits you to evaluate the quality of multiply stained sections prior to using the confocal.

If you first obtain optimal quality images on the standard fluorescent scope using NIH Image, you can then combine the images to produce a useful RGB Stack. This can be tentatively evaluated using the NIH Image function RGB to Indexed Color. The Stack can also be saved and produces a high quality 24-bit RGB image that can be further modified in Adobe Photoshop.

A)	Collect the image that you wish to be in the red slice of the Stack, using on-chip integration. Adjust the gain and black level to obtain the widest range of values, as indicated in a histogram of the image.

B)	Run the macro provided "Make Stack from Current Image". This will take the current selected image, make a new Stack with three slices, for the R, G and B planes, and paste the selected image into the "red" plane.

C)	Now collect a second image using the on-chip integration. Copy it to the buffer and paste it into the second slice ("green" plane).

D)	If you have a triple labeled section, collect this with the on-chip integration routine, place it in the third slice ("blue" plane).

E)	You can now use the RGB to Indexed Color operation under the Stacks menu to generate an indexed color image showing the double or triple labeled result.

F)	Save the file as an RGB TIFF file. It can now be opened with Adobe Photoshop for better color images.

The steps in the above procedure can be further automated with a macro.

III.	OPENING CONFOCAL IMAGES IN NIH IMAGE

The most commonly used CLSMs are made by BioRAD, Zeiss, and Leica instruments. The software/hardware used to generate images in these instruments stores the files in a DOS based file format. Thus the user is confronted with two initial tasks: transferring the file from a DOS based computer to a Macintosh and converting the file from its original file structure to an NIH Image file.

1.	Transferring Files From A DOS, DOS/Windows Or OS/2 Based Computer To A Macintosh

I suggest that you initially collect your images to the hard disk connected to your system. This will speed up data collection of extended Z-series. At the end of each work-session you can transfer all the files to your Macintosh using one of three methods: A) Sneaker net, B) Local area network, or C) FTP via Internet.

A)	Sneaker Net

Transfer the files from the storage location on the original disk to a Zip, Jaz, Syquest, magneto-optical or similar high capacity medium. These disks can be formatted as a DOS disk, and data directly transferred from BioRAD, Zeiss and Leica CLSM computers to these media. The Macintosh PC Exchange provided by Macintosh in System 7.1.2 and higher allows you to directly read PC-formatted floppies, Zip disks, and others. I don't recommend using Zip drives for original data collection as they are relatively slow, with access times of ca. 30 msec. The Jaz and Syquest drives have access/write times approximately equal to that of hard disks. For further comments and cautions, see below.

These disks can be manually carried to your Macintosh, which is presumably equipped with a similar drive. This is commonly referred to as a "Sneaker net". The major limitation of this method is that it demands that your Macintosh be able to read the various formats of different media. This is not always a valid assumption and relies upon your ability to find the correct drivers for different operating systems for the different types of media.

My own preference is for a disk of reasonable cost that can hold a typical workday's worth of data. The Zip drive, with storage capacity of ca. 100 MB and at a unit price of about $12-15, though relatively slow, meets both these criteria.

I don't recommend the use of floppy disks. They are very slow and cumbersome if you have many files, and requires that you have many pre-formatted disks. If you have an extended Z-series on a BioRAD CLSM, the file may be 7-25 MB, and is not easily transported via floppy disks.

Magneto-optical disks have a large capacity, and are long-lived. However, each drive manufacturer seems to have a different formatting scheme. I recommend them for archiving your files, but not for transferring them from the CLSM to the Macintosh. We have had endless aggravation with various incompatibilities.

B)	Local Area Network (LAN)

An efficient method of transferring files is to have both the DOS and Macintosh based machines on a common Ethernet network, with a common local server. If you have an Office of Computer Services (or something of similar ilk), they can provide various options for LANs between a Macintosh and a PC. An alternate means is to use MacLAN 5.5 for your PC MacLAN (Miramar Systems, Inc., Santa Barbara, CA) operates within Windows 3.1 and Windows 95. MacLAN allows your Macintosh "Chooser" to see the PC as another client on an AppleTalk zone, and your PC to act as if it were another Macintosh on your AppleTalk network. Assuming that you are familiar with using a Macintosh on a Ethernet network, this allows simple file transfers at high speed. We have found MacLAN to be a good idea, but seems to be plagued with bugs and problems in reliability and speed of transfer of files.

C)	FTP Via Internet

Transfer between computers that do not share a common server or network can be accomplished using Internet protocols such as FTP (file transfer protocol). In our experience, the most efficient program for doing this on the Macintosh is Fetch 2.1.2 (or higher), a public domain freeware program available on the zippy.nimh.nih.gov server. Fetch 2.1.2 will facilitate all aspects of the transfer from a PC/DOS machine. If you do not know how to set up your PC as an FTP site, contact your local computer center for assistance. The advantage of FTP using Fetch 2.1.1 on the Macintosh is that it doesn't seem to care what kind of operating system is working on the remote host. Fetch may be one of the most bomb-proof programs I have worked with--and it's free!

When performing an FTP, make sure that you perform the transfer in binary format, or the files will be unusable.

Do not erase the original files until you are certain that you have obtained a successful transfer and conversion. We recommend that you always save copies of the original files in their original native (DOS or OS/2) format.

A potential drawback to both LAN and FTP transfers is that the host computer containing the original files must be continuously available for the transfer. If someone decides to turn off the computer, or wants to use it for further data collection before you get back to your computer to effect the transfer, you may be cut off. In addition, if you are mainly a Macintosh user, you may find that setting up LANs and FTP links from the PC side is not as simple as on the Macintosh.

D)	Zip Drives

The new Zip drives from Iomega have inexpensive removable media that hold 100 MB/disk. This is usually sufficient for a single day's CLSM collection. A major advantage of the Zip drive is that it is also compatible with the IBM OS/2. I have used this without difficulty with the Leica/Windows 3.1 version, BioRAD Comos for DOS, and LaserSharp 1024 running under IBM's OS/2. The advantage of the Iomega Zip drive is that the drive is small, relatively inexpensive, and is supplied in two different hardware versions that can be connected to either the standard SCSI port on the Macintosh and many PCs, or to the parallel port on PCs. Iomega provides drivers for use with MacOS, Windows 3.1, Windows 95, Windows NT and IBM's OS/2. The disks, if formatted for the PC, are also suitable for OS/2, and can be read on a Macintosh using the standard PC Exchange provided in the Macintosh operating system.

1.	Few Words Of Caution About Using Zip Disks

The Iomega manual does not imply that a Macintosh disk can be read on a PC, so be careful about your formatting if you plan to use the disk on both a Macintosh and a PC. The drivers for the Iomega Zip do not operate as smoothly as they should. If the drive is formatted for use with a PC, you may have a problem displaying the root directory on your Macintosh. The disk seems to open, but it seems as if the root directory opens in a window that is "off-the screen" rather than on the monitor. I have narrowed down the problem to one of the Extensions I use with the MacOS 7.5.5, but I don't know exactly which one it is (I suspect one of the "Now Utility" Extensions). A kludgey work-around is to make a master folder on the Zip drive while still on the PC, and then put the sub directory folders inside that. Then, when using the Macintosh, if the root directory opens off the screen, select the Zip drive, and Open it. Then perform a Select All operation and Open again. This will usually open a window on the screen containing the sub directory.

If you open a file on your a PC-formatted Zip disk on the Macintosh, and then Save it, you will have difficulty using the disk again on a PC. Iomega has not been helpful about the issue of cross compatibility of their drivers between PC and MacOS.

In view of the problems I have encountered with the Zip drives mentioned above, my current practice is to transfer the contents of the PC-formatted Zip drive to a magneto-optical with Macintosh formatting. If I want to use the files on a Macintosh (e.g., at home), I transfer the files to a Zip disk that has been formatted for the Macintosh.

The archival stability of Zip drives is still unknown. I strongly recommend that you store the original data files on archival media.

E)	Magneto-Optical Disks For Archival Storage

The simplest means of storing large quantities of data is to store the file on a removable disk medium of large capacity that can be read by DOS, Macintosh and OS/2 systems. Since confocal images generate large data files, most systems have magneto-optical (MO) disks of 500 MB to 1.3 GB. A popular medium for this is the 1.2 GB MO disk, sold by Sony, Verbatim and others. These have the advantage of large capacity and have an archival life of at least thirty years. The most common drives are the Tahiti Max Optix 3, Pinnacle Sierra, various Sony, Ricoh, HP and NEC units. The disks are frequently (though not always) interchangeable. Many BioRAD confocal

scopes were supplied with Panasonic MO drives. These use a proprietary disk that cannot be read by the drives of the previously mentioned manufacturers. When confronted with a Panasonic drive, use a LAN or FTP to transfer your files.

Until recently, these were the most common means of primary data storage for confocal microscopy. The price of the disks have dropped recently, and are now available for about $50 for a 1.3 GB disk. This is about half the price of a Jaz drive (also 1 GB), and though slower than the Jaz drives, is less prone to data loss.

The popularity of magneto-optical disks has declined recently due to the widespread availability and low price of Zip drives. Many people have had problems when attempting to read standard 1.2 GB disks that were originally formatted on a DOS based Pinnacle or Tahiti drive. This is particularly severe if you are using OS/2 with LaserSharp 1024. A significant problem with magneto-optical disks is the lack of universal driver standards. Use "Multi-Driver" on your Macintosh to facilitate reading DOS-formatted optical disks. "Multi-Driver" is a Macintosh Control Panel included in a software package sold by PC Access.

The BioRAD OS/2 optical driver cannot read Macintosh-formatted magneto-optical disks at this time. I no longer recommend using magneto-optical disks to move files from a PC platform (Windows or OS/2) to a Macintosh. Go with a Sneaker net using Zip drives, LAN or FTP. Magneto-optical disks, however, remain an ideal medium for archival storage of large files.

2.	Opening BioRAD Files in NIH Image

BioRAD files consist of a 76-byte header that defines the size of the image in width and height, states if it is a Z-series, and, following the image data, provides information on parameters during collection of data, magnification scale and notes. The file structure of the new software, LaserSharp 1024, is similar to that of the earlier versions of Comos, with only minor exceptions pertaining to the footer at the end of the file that contains information about the parameters used during data collection.

Individual BioRAD files can be opened using the Import function of NIH Image. However, this requires that the user know the width and height of the individual file (Most frequently 768x512 in BioRAD files from an MRC600; 512x512 or 1024x1024 in LaserSharp 1024 for an MRC1000; 512x512 in Leica files). Unfortunately, I cannot figure out how to transfer a merged 8-bit BioRAD image to a Macintosh. The associated LUT (look-up table) is apparently stored in a manner or location that eludes me. If someone has solved this problem, please let me know. The Import function does not permit importing a Z-series. The macro for importing BioRAD files will correctly import a Z-series.

The macro "Import BioRAD MRC Z Series ", contained in the associated file "Confocal Macros", enables opening of either single sections or a Z-series. This works equally well on files generated with Comos and those generated with the recently released software, LaserSharp 1024. LaserSharp 1024 operates under OS/2. These files can be opened using the same macro employed to open the DOS-formatted files. LaserSharp 1024 stores each separate color plane of an RGB Z-series as a separate file. All three image planes of a single RGB file can be optionally stored as a single file.

After loading the BioRAD macro, you can "run" it by selecting it from the Special menu. The dialog will ask you for the "Starting Slice" of the Z-series. The default is 1.00. Accept that value for the moment. A BioRAD file containing only a single image will appear on the screen with the title of the original file (e.g., Axons03.PIC). I suggest that you immediately Save the file, prior to making any modifications, (e.g., “Axons03.PIC.Img”). This macro will load as many sections as possible, dependent upon memory.

If the number of Z-sections exceeds the memory allocated to NIH Image, the macro will stop running. You can Save the first group of sections in a separate file. Close the file, and again run the macro "Import BioRAD MRC Z Series", but when prompted for the "Starting Slice", enter the number of the last Z-section displayed. If the original Z-series is very large, you may have to do this several times. You can solve this problem either by buying lots of RAM, or store Z-series in smaller increments. This problem is particularly notable if you have collected images at 1024x1024 as well as collecting an extended Z-series. (e.g., 70 sections at 1024x1024 results in a file of ca. 74 MB.)

The associated calibration and notes of a single image file will be placed in a separate NIH Image text window. If you wish to include this information on the image that you saved as an NIH Image file, you must Paste it from the text window to the image window. If your notes contain information about pixel size and dimension of the image, use that to calibrate the scale of the image. (See "Set Scale" in NIH Image Manual). Once you have calibrated the image, Save the file once again. The current version of the BioRAD macro does not properly read the calibration or notes file attached to a Z-series file.

BioRAD Z-series can be stored as a single large file, or as a series of individual sections. The former configuration has the advantage that all related files are stored in a single locus. However, it also means that such files may be huge. Using the above macro, a BioRAD Z-series will be opened into a new Stack. Save this in a similar manner suggested above for a single image file.

If this is a Z-series, use Stacks menu, and select Options to enter the step size of the motorized focus used when originally recording the Z-series. The magnification and slice spacing are stored with the original BioRAD file.

You may find it useful to Add a Slice to the beginning of a Stack to record specific comments about the series, what the data represents, and an image of Z-projection to demonstrate the contents of the file.

Save the original BioRAD *.PIC files, just in case. You can move all the original files into a separate sub directory to reduce clutter.

A)	BioRAD Split Screen Images

Simultaneous collection of double labeled sections on the BioRAD can be displayed and saved on a "split screen" image, with the two images from PMT 1 and 2 displayed "side by side." This has two advantages:

1.	You can directly compare the location of different antigens simultaneously

2.	When collecting a Z-series with simultaneous split screen, you can be assured that the two images will be in the same focal plane.

The Z-motor drive on the BioRAD has not proven as reliable as it should be, and is prone to slippage due to poor mechanical linkage. If the Z-series for two or more antigens is collected sequentially, you may not be able to rely on the accuracy of Z-position to return the section to the same vertical location in subsequent scans. The inaccuracy may also be due to the current use of sloppy focus mechanisms on most modern microscopes. The modern co-axial planetary gears do not reliably return the stage to the starting position when restored to the same position on the fine focus knob.

The macro described in the previous section will open and display the split screen image. If you want to separate the two halves of the screen into separate images, use the macro "Merge BioRAD Split". This will convert the original split image into a three slice Stack (left image, right image and one blank black slice), then produce a merged color image. If you want to interchange the red and green planes, Select the window containing the Stack of three images, and run the macro "Swap Red_Green".

The Stack of three slices (left, right and blank), can then be saved as a 24-bit image in Adobe Photoshop 3.0 format, in the following manner:

1. If you have run the preceding macro, and have generated an RGB image in NIH Image, NIH Image will save the file as an RGB-TIFF file. This format is accepted as an Adobe Photoshop 3.0 RGB image.

2. Save this file with a modified title in order to preserve your original file.

3. Open the file using Adobe Photoshop 3.0.

The quality of the RGB image produced within NIH Image is an 8-bit indexed image, is of marginal quality, and should be considered to be a crude "proof" image. You may find that the image can be greatly improved by changing contrast or brightness of the individual slices. Before you do this, make sure that you have saved a copy of the unmodified Stack. In general, you will find that such modifications are best accomplished in Adobe Photoshop 3.0.

B)	Merging Split Screen Z-Series

The macro that operates on a Z-series of a split screen multi-labeled section will:

1)	Place the left image in one Stack, and the right image in a second Stack.

2)	Color Merge the two Stacks, and that will allow you to Animate the Stacks. This is very helpful when you want to track individual processes in a complex field through a lengthy Z-series.

(I am also planning to add a macro that will take a third Stack and place it in the blue plane of an RGB Stack.)

3.	Leica Files

Leica file format is a TIFF format. The Leica operating system is a peculiar hybrid using a VME bus with a 68040 CPU operating under OS/9, but uses a Windows front-end that generates DOS types of files. Leica has announced the release of a new software/operating system for their confocal microscope, based on Windows NT. The format of the images, however, will reportedly be unchanged. The original Leica files (and the new Windows NT files) can be directly read by a Macintosh. They will show

up on your Macintosh desktop as PC files, with various optional icons, depending on how you set your system parameters in the Control Panel, PC Exchange. Each separate image file is accompanied by an "info.dat" file. Z-series are stored as single files with sequential numbers but with only a single info.dat file for the whole series. In order to reduce confusion in handling these Z-series with large numbers of files, move each Z-series set into a separate folder on your Macintosh.

If you are running NIH Image, you can Open these files directly, without having to Import them. However, you will not be able to double-click on the files to open them in NIH Image. In order to do that, pursue the following procedure:

A)	Obtain a copy of “CTC 1.4” (or later) to change file type and creator. CTC 1.4 is available from the NIH Image server at zippy.nimh.nih.gov.

B)	Select all the Leica graphics files in each folder (not the information.dat files) and drag them on top of the icon for CTC.

C)	Using the resulting dialog box, change the "New Creator" to Imag, and the "New Type" to TIFF (creator and type are case sensitive, so enter exactly as spelled).

All the files will now have an NIH Image icon, and will be treated as NIH Image files by the program.

The info.dat files should be changed to "New Creator" Imag and "New Type" to TEXT. This file contains important information including scaling factors, size of image, size of pixel, and spacing in a Z-series.

To open a Leica Z-series Stack:

A)	Close all other NIH Image windows at this time. Make sure that you have placed all the related Leica image files into a separate folder.

B)	Select Open from the File menu. In the resulting dialog box, go to the desired folder and click "Open All", then Open. This will open all the files in rapid sequence.

C)	Select Windows to Stack from the Stacks menu. Make sure that you set the magnification/calibration scales and slice spacing.

D)	Save the resulting Stack with a new name.

4.	Zeiss Files

Zeiss files saved using the newer version of software with the LSM 310 and 410 are “standard” *.TIF files. Use the same procedure outlined above for the Leica files--i.e., use CTC to change "New Creator" to Imag and "New Type" to TIFF. The files can then be directly opened by NIH Image. Make sure the LUT is correctly set. I have often found the Zeiss LUT map to be inverted, with a negative slope, and an inverted image. Click on the icon in the lower left corner of the mapping window to obtain the correct black/white relationship.

5.	Molecular Dynamics/Sarastro

Jay Hirsh kindly provided information on the format of Molecular Dynamics files. The data files are simple files without a header or footer. Each image of a Z-series is stored

as a separate file. Information about a single image, or about a Z-series is stored in a separate text file. The text file can be directly opened by NIH Image.

A)	Place all associated images of a single Z-series in a single folder. Make sure that they are correctly numbered to reflect their order of collection (e.g., 008.ext, 009.ext, 010.ext, not 8, 9, 10).

B)	Select Import from Stacks menu. The user must know the image format (512x512 or 1024x1024) in order to use this correctly. The offset value is 0.

C)	Check boxes "Open All" and "Invert".

D)	Click Open. This will open all the files to the screen.

E)	Select Windows to Stack from the Stack menu.

F)	Save the Stack with a name of your choosing.

6.	Noran Confocal Files

There is a version of NIH Image that is used for direct data collection on the Noran. The resultant files are obviously compatible with NIH Image. The Noran is also sold with a version of Image-1, a DOS based program. A version of this program is now also available for DOS/Windows. I have no information about the file format used by Image-1.

IV.	BASIC IMAGE MANIPULATIONS

NIH Image provides the user with an extensive range of image processing tools. You should become fully conversant with all these tools, including how to obtain a histogram, how to interpret the histogram, manipulating the look-up table (LUT), substituting a pseudocolor LUT in place of the standard gray scale LUT and inverting the LUT.

Keep in mind that most changes will only be made to the "display" buffer not to the "file" buffer. Thus, if you examine a histogram, then modify the brightness and/or contrast, the histogram will not be changed. If you now Apply LUT to the file buffer, the new histogram will reflect those changes. The original file, stored on your disk, will not be altered by this operation unless you now Save this modified file. The reader should use one of their own sample CLSM files to familiarize themselves with  operations of NIH Image. All the operations in this section are fully dealt with in the NIH Image manual, About NIH Image.

You now can open one of the files that you imported to the Macintosh. Until you are experienced with the program, make a backup copy of the file, and only work on the copy, not the original file. Using the Save As... function of NIH Image, change the name of the file so that you don’t mistakenly modify your original data file.

1.	Evaluating The Quality Of Your Original CLSM Image

As stated previously, the quality of the image that you obtain from the following manipulations will be directly dependent upon the quality of the image that you start with. Thus, if you start with a lousy image, you may be able to make it look presentable, but it is still going to be a lousy image.

A)	Making sure that your original CLSM image uses the full range of 8-bit values (0-255)

The most common flaw with confocal images, as with video images in general, is that the user has not utilized the full 8-bit range of gray values. Many of the manipulations that you perform on the original CLSM may make your image appear to be satisfactory. However, many data files do not contain a full dynamic range of values (0-255), and the image was massaged by artificially spreading a narrow range of values (e.g., 50-120) by modifying the LUT.

In order to develop an appreciation for the gray scale content of your images in NIH Image, examine a histogram of the image.

1.	Make sure that the image window is active by clicking anywhere within the window.

2.	Now hit the Command+H key combination.

This will open a separate window with a histogram showing the distribution of gray scale values (from 0-255) on the X-axis, and the number of pixels for each gray scale value on the Y-axis. Ideally, the histogram should be spread out over the full range of values. The most common form of error is to have an extremely contrasty image with all the values clustered at one extreme end of the histogram.

Learn to use the adjustments to black level and gain on the CLSM to optimize the spread of gray values. If you have to trade off between using the "+1 to +3 LUT" position on the BioRAD versus a brighter setting of the laser, choose the brighter laser setting. It will burn out your specimen faster, but give you a better signal to noise ratio and a wider dynamic range in gray scale value. The "Photon Counting" mode, or the "Accumulate" mode on the BioRAD will often provide an image with the best dynamic range. On the BioRAD, I prefer to use the "Accumulate” mode with "Slow Scan" for best results. It is not the purpose of this manual to teach the use of CLSMs, but I only wish to emphasize the importance of the quality of the original image.

B)	Avoiding High Contrast Images

If you get into the habit of checking the histogram of your images as you collect them on the CLSM, you will improve the quality of the original images, and find less need to twiddle with the LUT values.

C)	Avoiding Noisy Images

If possible, collect images using the "F1" (Slow) setting (on the BioRAD) with a Kalman setting of at least 3, or "Accumulate" mode. When using "Accumulate mode", you can use a less intense laser source, and manually accumulate until the image appears satisfactory.

2.	Editing Image

A)	Cropping Images, Erasing, Superimposing Text, Scale Bars, Rotating And Scaling Images

NIH Image provides a wide range of tools for editing your image. These are described in NIH Image manual About NIH Image.

1.	Notes On Scaling

You may find that you want to enlarge a selected portion of an image to emphasize a particular point. For optimal images, do not use the magnifying tool for enlarging images prior to printing. Use the Scale and Rotate function of the Edit menu. Select the option "Bilinear Interpolation" method, rather than "Nearest Neighbor." The resultant image will appear much smoother, with less "pixellation". The "Bilinear Interpolation" function will operate correctly on an indexed color image. In order to avoid pixellation in this instance, move the components of the RGB Stack to individual windows, Scale/Bilinear, then move them back to an RGB Stack and convert to 8-bit color. The macro "Crop and Scale-Smooth" will simplify this operation on Stacks.

2.	Caution On Scale And Rotation Of Images

The original confocal image may not be oriented at an optimal angle for the final publication. You can easily rotate digital images in order to align them for your final illustration. However, this results in potentially severe image degradation associated with digital rotation. This will not occur if your you rotate the image in 90° increments. However, any other rotation will result in progressive degradation due to smoothing functions (interpolation). You can best demonstrate this by taking an image containing many fine, threadlike profiles, and rotate it in 10° increments through 360°. Test this using both linear and bicubic interpolation. Place the resulting images into a Stack and Animate the Stack. You can now directly observe the increasing blurring of the fine processes consequent to the rotation of the image. In order to minimize this effect, I suggest that you determine the final angle of rotation desired and make only a single rotation of your original image. If the angle selected is unsatisfactory, discard the result and start again.

3.	Using LUTs

The following operations are commonly used to enhance the image. These are standard operations on all confocal scopes, and are based on methods that are widely used for manipulating digital images. See appropriate section of the NIH Image manual.

A)	Modifying Brightness And Contrast

In the simplest operation involving look-up tables, you may choose to emphasize a selected range of index values (gray scale values), and minimize other values. The choice may be based on your interpretation of the information content of the image. Learning how to interpret the histogram, and modifying brightness and contrast are essential skills in image processing.

B)	Linear And Non-Linear LUTs, Including Custom LUTs

The standard LUT shown in the mapping window (lower left side of your screen) is a linear LUT. You can modify the brightness and contrast values by either dragging the Brightness and Contrast slider buttons, or by directly dragging the plotted line in the mapping window.

The "Confocal Macros" file included with this manual provides a number of alternate, non-linear LUTs based on sampling the values within your display buffer. These include Log, Parabolic, Square, Square Root and Gamma transforms. Play with each

of these and observe the effects on your images. You will probably find the Gamma transforms most useful when used with a setting of 1.5 to 2.0. Many of the resultant changes produce results similar to those obtained on the BioRAD with +1, +2 and +3 LUT settings. When you open BioRAD files with NIH Image, the results may not match what you saw on the BioRAD. Most commonly, the image will appear much darker, and lacking the detail you so clearly remember seeing on the screen when collecting the original image. NIH Image has not corrupted your files. This is most frequently due to the fact that the image you viewed on the confocal microscope may have had a non-linear output LUT attached to it. The original data is imported intact, but the output LUT is not attached to the new NIH Image file, and a new linear LUT is appended to the file. See the NIH Image manual for instruction on how to modify the LUT.

C)	Enhance Contrast Operator In NIH Image

The Enhance menu item in NIH Image has a specific function entitled Enhance Contrast. This produces a custom linear LUT effect that is as good as any result I can obtain in my attempts to manually modify the LUT curves. However, you may find that this operation produces an excessively contrasty image, with too steep a slope in the LUT mapping window. If so, then manually change the LUT in the mapping window to give a slope of the LUT about halfway between the original value and that produced by Enhance Contrast. If that is satisfactory, then Apply LUT (Command+L). If you feel that you want still more contrast in the image, repeat the above sequence.

In order to appreciate the effects of this operation on the original image, examine a histogram of the image before and after this procedure.

As in all alterations to the LUT, the Enhance Contrast operation only modifies the display buffer, not the image in the file buffer or the original file on the disk. In order to alter the file buffer, you must perform Apply LUT. This will not alter your original disk file. If you wish to do so, Save the file. Once you have done that, you cannot go back to the previous image. You may, therefore, prefer to save the modified file under an alternate name.

Enhance Contrast is very different from the Equalize operation, which is likely to result in excessively splotchy images. Compare the effect of each of these operations on the appearance of the mapping, histogram windows and LUT.

D)	Thresholding And Density Slicing

See NIH Image manual.

E)	Pseudocoloring Images

Pseudocoloring confocal images assists the viewer in displaying double labeled sections, detecting major differences in concentration, or changing concentrations, as in calcium ratio imaging. See appropriate menu item, and NIH Image manual for use of this operation. All the color applications described in this manual, although they may resemble the original colored fluorescent image, are pseudocolor.

F)	Exporting To Adobe Photoshop

Recent versions of NIH Image save single files and RGB files in a format that can be directly opened by Adobe Photoshop 3.0. This will not work with Z-series.

4.	Enhancing/Filtering Image

NIH Image provides a wide range of tools for enhancing and filtering images, including filters to sharpen, smooth, shadow and detect edges. You can write your own kernels, run median filters, and Sobel operators. Some of these operations are multistage operations and alter both the display buffer and the file buffer (but not the disk file), and cannot be Undone. You will have to reopen the original file to restore the image.

For further details, see NIH Image manual.

5.	Quantitative Measurements

NIH Image was originally developed for quantitative measurements. There are many useful functions described in the About NIH Image manual. Various measuring functions include: length, area, density, particle counting, and profile of a Line. About NIH Image will provide guidance in the use of these operations.

6.	NIH Image Macro Language

One of the most powerful features of NIH Image is the macro language. This allows the user to write simple Pascal-like scripts. The distribution kit provided from the FTP site contains many sample macros that can be readily modified to your particular needs.

V.	Advanced topics

1.	Merging Pairs Of Double Labeled Sections

One of the most valuable and commonly used features of confocal microscopy is the ease of obtaining images of double and triple labeled histological sections. There are two major obstacles that may cause difficulties when attempting to merge two images:

A) The contrast and brightness values of one section may be markedly different from the other(s). This is best dealt with by careful evaluation of the images at the time of original data collection, and modifying your means of collection.

B) The two series of images of such a pair may not be in perfect register with each other. This may be due to various factors, including misalignment of the PMTs, the mirrors, and filter blocks. Most of the errors appear to occur in translation (X- and Y-axes) and not in rotation. A shift of a few pixels may not be noticed, but occasionally the error results in a marked shift from one color plane to the second. Simple translation errors can be corrected by shifting the images one or more pixels at a time. Rotational errors are more difficult to correct, take longer, and frequently result in image warping. If you find that you have marked rotational errors, the service personnel from the manufacturer of your confocal microscope should deal with this.

NIH Image provides an alignment operation, "Register". This can be used on sections in a Stack (see NIH Image manual).

A)	NIH Image And Adobe Photoshop

Two or three gray scale images of different fluorophores can be combined into a single colored image, with each fluorophore represented by a different color. The color chosen to represent each fluorophore is arbitrary, and can differ from the original one. There are two different programs that can be used successfully for this purpose, NIH Image and Adobe Photoshop. There are benefits and disadvantages in the use of each program. Adobe Photoshop is relatively expensive, but a superb commercial program. It supports 24-bit images and allows almost instantaneous adjustment of the individual red, green and blue planes of a merged image. NIH Image produces an 8-bit custom palette of the merged image. It takes 5-10 seconds to produce this image on a Macintosh IIfx, and is correspondingly faster on a Quadra 950 or a PowerPC. Although this is an 8-bit color image, the result is often satisfactory, and occasionally even comparable to that obtainable with Adobe Photoshop. However, NIH Image may produce excessive dithering of the resultant image, and you cannot make small adjustments in brightness or contrast of the final color image obtained with NIH Image. Instead, you have to go back to the original gray scale images, modify them, and then once again Merge the images.

Since each "Indexed Color" image produced with NIH Image has its own unique LUT, you cannot directly do a side-by-side comparison of two different color images if you merged using "Custom Colors", as the color values shift markedly as you change windows. If you selected "System Colors", the quality of the color is more limited, but you will be able to compare results with other windows merged using the same system LUT. This is a major disadvantage in relying exclusively on NIH Image. However, for routine operations, NIH Image is satisfactory.

In comparison, Adobe Photoshop, using a full 24-bit window, allows you to compare multiple colored images simultaneously on a single screen. Adobe Photoshop also provides an excellent range of filters and convolutions. The more immediate advantages of NIH Image are manifest in measurement capability, generating Stacks, Z-series projections, and 3D-projections and rotations. Adobe Photoshop does not provide such facilities.

Until recently (prior to version 1.56), NIH Image could only run under an 8-bit monitor setting. If you wanted to shift back and forth from NIH Image to Adobe Photoshop, you had to reset the monitor to 24-bit. Beginning with version 1.56, NIH Image operates satisfactorily with the monitor set to 24-bits.

NIH Image version 1.59 now directly allows the user to save an RGB stack of three sections in a format that can directly be read by Adobe Photoshop 3.0. If the file is modified in Adobe Photoshop, then saved as an Adobe Photoshop TIFF file, it can be re-opened by NIH Image as a three slice stack. However, if you add additional Layers or Channels in Adobe Photoshop, it forces you to Save the file in Adobe Photoshop format. This cannot be read by NIH Image. An NIH Image Z-series containing Merged color slices cannot be read by Adobe Photoshop.

B)	Double Labeled Sections: Building A Stack (Best Done Using A Macro)

1.	Color Merge In NIH Image

Close other windows. Once you become more adept at handling and opening new Stacks (using a macro), it will not be necessary to close other windows.

Two color plate: If you wish to combine only two plates, open the two files.

Use the macro "Color Merge Two Images" contained in the sample "Confocal Macros" that is provided with this manual. If you examine the sequence of commands in the macro, you will find the following operations:

A)	Open a fresh stack

B)	Paste the "red" image into the first slice

C)	Add Slice to add an additional slice

D)	Paste the "green" image into the second slice

E)	Add Slice to add a black (empty) third slice

F)	Merge from RGB to 8-bit color using a custom LUT

G)	Open a new window containing the merged color image

Alternately, you can do this operation manually to familiarize yourself with the procedure. The first file opened will become the red plane, the second file the green plane. Under menu item Stack choose Windows to Stack. This will put the two plates into a stack labeled "RGB". Any further operations will not alter your original data files, so you can always go back and start again.

Save a copy of the NIH Image Stack. The following section describes how to use this Stack with Adobe Photoshop 3.0.

If you want to alter the contrast or brightness of either of the color planes, you have to modify either the original images or those in the RGB stack. I suggest that you limit your initial attempts to the slices in the RGB stack. When you change the Brightness or Contrast of any single image, you must then Apply LUT (Enhance menu) to that single image. Do not use the Apply LUT to "Stack" macro as this will modify all the images in the Stack.

From the Stacks menu select RGB to 8-bit Color. In the resulting dialog box, select "Custom Colors". This will generate an "Indexed Color" window from the Stack, but will not alter the stack itself.

Since each original file may have a different distribution of values (see histogram), the saturation of each color plane may differ markedly. Be creative, try different ways of doing it. (e.g., use Enhance Contrast operation on one slice at a time). Try non-linear LUTs, various filters, etc.

The resulting "Indexed Color" image can be saved with a unique name.

Once you have mastered this sequence, you will have a clearer understanding of the operation of the macro "Color Merge Two Images" contained in the sample "Confocal Macros" that is provided with this manual.

If you now Save the RGB stack (RGB TIFF) in NIH Image, the resulting file can be viewed with Adobe Photoshop as a 24-bit file.

2.	Merging Files With Adobe Photoshop

The recently released version of Adobe Photoshop 3.0.1 has excellent layer management.

Version 1.58/1.59 of NIH Image now permits Stacks to be directly saved in Adobe Photoshop 3.0 TIFF format.

A)	The stack must consist of 3 slices.

B)	Before saving the stack, open Slice Info in the Stacks menu. Confirm that RGB is selected.

C)	Save the file. It will write the file as an RGB TIFF file.

You can then directly open this NIH Image Stack in Adobe Photoshop 3.0.

An alternate means of converting the stack to an RGB TIFF format is to call the RGB to Color function. This will automatically change the "Slice Info" window to RGB. The resulting "Indexed Color" image will give an approximate (8-bit) preview of the results to be obtained with Adobe Photoshop (24-bit). Save this file.

You will now be able to use the various Adobe Photoshop tools to modify the resulting 24-bit image. You should explore the use of the Levels command (Command+L), as well as the Brightness/Contrast command (Command+B).

If you have a quadruple (or more) labeled section (e.g., 512x512, four fluorophores, or three fluorophores and a Nomarski/DIC image), and have stored them in a four slice stack in NIH Image, you can open them in Adobe Photoshop using the method described below.

A)	From the Adobe Photoshop File menu, select Open.

B)	Using the Adobe Photoshop Open dialog, go to the desired directory containing the NIH Image Stack file of interest.

C)	Using the same dialog box, pull down list at the bottom of dialog box, choose Raw file format.

D)	This will show all files in the chosen folder.

E)	Select the NIH Image Stack containing the four slices. You have to know the dimensions of the slices.

F)	The resulting dialog box requests that you fill in:

1.	Width (e.g., 512 pixels)

2.	Height (e.g., 512 pixels)

3.	Number of channels (e.g., 4, for the number of slices in the stack)

4.	Select "not interleaved"

5.	Offset =768 (3 X 256 = 768)

G)	The original NIH Image stack will now open as a four channel color stack in Adobe Photoshop. However, the colors in the image are inverted. Invert (Command+I) the image to obtain an appropriately colored image.

H)	Save the new Photoshop image using the Save As... dialog box. Do not save as a Raw image. Select TIFF format, and choose a new name for the file, otherwise the simple Save command will overwrite the original NIH Image file and store the data in the Raw format.

You can modify each color channel separately using several different methods, as described in the Adobe Photoshop 3.0 manual.

The quality of the resulting color image will generally be much better than that provided by NIH Image, as it can utilize a full 24-bit look-up table, and does not require dithering of the image.

The Adobe Photoshop manual and tutorial will provide further guidance in modifying the images.

3.	Preferred Method

It is obviously easier to use the new procedure provided in NIH Image version 1.58. The first slice of the NIH Image Stack forms the red layer, green the second layer, and the third slice forms the blue layer. The gamma, brightness and contrast of each layer can be individually modified with immediately evident results. For further details, see the Adobe Photoshop 3.0 manual.

C)	Compare Results Of NIH Image 8-Bit Merge With Photoshop 24-Bit Merge

The quality of the color image (assuming your monitor is set for 24-bit color) is generally much better in Adobe Photoshop than the optimized 8-bit image obtained in NIH Image.

D)	Adjusting Color Contrast On Sections In A Stack (See Macros)

Do the Enhance Contrast operation separately on each section in the stack. Failure to do so will result in the merging of both saturated and unsaturated images.

If you are dealing with a single RGB section, you will find that Adobe Photoshop provides much better tools for this.

E)	Merging A Double Labeled Pair Of Z-Series Using A Macro

1.	Open the two Stacks. The first stack will be assumed to be the "red" stack. The second stack will be the "green" stack.

2.	Close all other windows.

3.	Select the macro "Color Merge Two Stacks".

4.	Lean back and watch the fun.

When completed, Select the "Merged" stack and Animate it, or step through it with the "<" and ">" keys.

The drawback of this procedure is that the resulting images are 8-bits, not 24-bits. I cannot find a way to do this procedure with Adobe Photoshop.

2.	Projection Of A Z-Series And 3D Rotations

Many of the most valuable qualities of confocal Z-series are that the data can be manipulated to obtain "through views" of the stack, the stack can be resliced at various angles converting a transverse view into a sagittal or horizontal plane, it can be used to generate 3D images, stereo pairs, and other operations. Indeed, NIH Image is able to perform operations as well as much of the software provided by the manufacturers of confocal microscopes, and equivalent to many of the basic operations provided by very expensive image processing programs such as VoxelView, VolVis and VoxBlast running on a Silicon Graphics Workstation. For more elaborate operations using true Voxels, complex shading and ray-tracing in 24-bits, these latter programs are excellent choices. For the most common operations, however, NIH Image is quite adequate. Optimal performance on these tasks will be achieved using a PowerPC with sufficient memory to handle your typical data sets.

A)	Stepping Through A Z-Series Using Stacks

1.	Open a Z-series in a Stack.

Use < and > to step through the plates in Stack, one at a time. The reader is urged to read the section in the NIH Image manual regarding Stacks.

B)	Animating A Z-Series

1)	From the Special menu, select Video for oscillating movies. This will produce smooth back-and-forth motion of the Z-series animation.

2)	Command± to Animate a series. Number keys 1-9 control the speed of the animation.

C)	Projecting A Z-Series Onto A Single Plane

This is an extremely useful benefit of the confocal microscope. A stack of Z-series sections are all in optimal focus. If projected onto a single plane, all objects throughout the thickness of the imaged section will now be in sharp focus and spatial relationships may be more evident.

NIH Image version 1.58 and later provides macros to perform the Z-projection. A sample macro to accomplish this operation is included in the accompanying macro files.

1.	Selection Of Area For Z-Projection

A)	Choose an extended Z-series (e.g., more than 10-20 images) with well defined profiles of objects.

B)	Run the macro "Project Z Series".

You will gain a better understanding of this procedure if you examine the macro, and will find that it merely automates the following procedures:

C)	Use Project command

D)	Set:

| Slice Spacing (Pixels): | Set with your slice spacing |
| --- | --- |

| Initial Angle (0-359): | 0 |
| --- | --- |

| Total Rotation (0-360): | 0 |
| --- | --- |

| Rotation Angle Increment: | 0 |
| --- | --- |

| Lower Transparency Bound: | 0 |
| --- | --- |

| Upper Transparency Bound: | 100 |
| --- | --- |

| Surface Opacity (0-100): | 0 |
| --- | --- |

| Surface Depth-Cueing (0-100): | 100 |
| --- | --- |

| Interior Depth-Cueing (0-100): | 0 |
| --- | --- |

E)	Select "Minimize Window Size"

F)	Select "X-Axis"

G)	Select "Brightest Point"

After you develop a sense of familiarity with the program, play with different values. Start with the default values. Once you are more familiar with the Thresholding tool (Magic Wand), you will be able to use that to set the range of values desired for your Z-series.

This will generate a single Z-projection of all the plates in the Stack, viewed from directly in front of the stack.

Experiment with different view angles by using different values for the "Initial Angle", while leaving "Total Rotation" at 0.

If you want to make a quick and dirty stereo pair of images from a Z-series, then the values should be something such as:

| Initial Angle (0-359): | 356 |
| --- | --- |

| Total Rotation (0-360): | 8 |
| --- | --- |

| Rotation Angle Increment: | 8 |
| --- | --- |

Select the "Y-Axis" as your "Axis of Rotation". This will cause the resulting image to rotate left to right (i.e., around  the Y-axis). Choose "Brightest Point" as the "Projection Method".

This will produce a new Stack containing two images at an 8° increment. This is a common angle for stereo pairs. In order to visualize this, you can do two different procedures.

1.	From the Stacks menu, Animate the stack, and the image will rock back and forth around the Y-axis.

2. Make a Montage using this function from the Stacks menu. You can experiment with different values for the scaling factor. The smaller the scaling factor, the easier you will find it to fuse the images. Ideally, for people with limited experience in "Free Fusing" stereo pairs of images on the screen, matching points on the screen should be less than 50 mm apart. With further experience, you will be able to fuse larger images with matching points further apart.

2.	Optimizing Settings For Project Function Of Z-Series

A)	Make sure that "Invert Pixel Values" are not selected in the Edit/Preferences menu item. If selected, de-select this item. Then

Select File/Record Preferences. Quit NIH Image in order to implement the changes in preferences. If this item is "selected", you will be easily confused by the various values displayed in the information, mapping and histogram windows.

Remember that a value of 0 equals the brightest level, and a value of 255 equals absolute black. This is a feature of the Macintosh Operating system setting, and also conforms to the original software that preceded NIH Image. This software was developed to measure the density of dark regions. Thus, a dark area had a high concentration of a substance, and correspondingly, a "high" value. Regardless of what your intuition may tell you, just accept this as a fact of life, and avoid the confusion that comes with protest.

B)	There are a number of parameters in the Project item under the Stacks menu. Understanding the proper use of this item is critical to obtaining pleasing results in Z-projections and rotations.

C)	The first of several items include obvious settings regarding "Slice Spacing", "Initial Angle", "Total Rotation" and "Rotation Angle Increment". These values are all obvious and pertain to the sampling interval of the resulting projection.

The remaining items on the list, "Lower" and "Upper Transparency Bounds", "Surface Opacity", "Surface-Depth Cueing", and "Interior-Depth Cueing", however, are very confusing, and cannot be easily understood without some explanation.

One of the greatest difficulties in mastering these functions is because once you have selected specific settings, obtained an image, and then try other settings, your screen will be cluttered with images, and the particular parameters used to obtain them has been forgotten.

Start with the "Lower" and "Upper Transparency Bounds" in relationship to the stereo pairs you generated a few minutes ago.

When making a simple single view Z-projection, the slice spacing is irrelevant.  However, depending upon the setting of the "Lower" and "Upper Transparency Bounds", and "Interior Depth-Cueing", or the resulting image is likely to appear somewhat lacking in brightness.

"Lower Transparency Bound" must be set to 0. (Note that this value is expressed from 0-254). This value determines the cut-off point of displayed pixels. All pixel values below the setting selected will be discarded and modified to a darker value. Remember the whitest/brightest pixel on the Macintosh has a value of 0, the black pixel a value of 255. "Lower Transparency Bound" greater than 0 (e.g., n), will modify all pixels from zero to n  in your image and replace them with darker pixels. This results in dark gray to black holes in the midst of a bright object. The setting of this value is quite simple. The most effective way to appreciate the consequences of this choice is to look at histograms of the images produced with "Lower Transparency Bound" set to zero, and compare that with a histogram generated with a "Lower Transparency Bound" of perhaps 50. Make sure that your starting image has prominent objects with histogram brightness index values of 1-10 (out of a range of 0-255).

For the "Upper Transparency Bound", I suggest that you start with a value of 254. After using this for a while, play with other values.

The three following functions in the Project dialog box are expressed as values of 0-100%.

| Surface Opacity (0-100): | 0 |
| --- | --- |

| Surface Depth-Cueing (0-100): | 100 |
| --- | --- |

| Interior Depth-Cueing (0-100): | 0 |
| --- | --- |

NIH Image 1.61 sets "Interior Depth-Cueing" to a default of 50. This will produce "muddy" images. You must set this to 0.

D)	Reslicing The Z-Series Along Alternate Planes: (X-Z, Y-Z And Theta-Z)

One of the most useful features of NIH Image  is the ability to rapidly reslice a Z-series data stack along alternate planes. In addition to simple orthogonal planes (i.e., X-Z and Y-Z), you can reslice at any arbitrary angle between X and Y to generate a single Z-plane cross section. This is very useful for evaluating the extent of penetration of antibodies into tissue, evaluating cell spacing, etc.

The closer the images in the original Z-series, and the greater their number, the more natural appearing the final results. In order to provide a seemingly continuous image in the resliced plane, the software interpolates gray values. For improved quality of resliced images, your original Z-series should be as close as possible (e.g., 0.25 µm).

Have you set the calibration for magnification and slice spacing, as described above? (Also see NIH Image manual).

Your original data file may contain the needed information about slice spacing. In many of our files, the step size was 0.38 µm (on a BioRAD), and the pixel size 0.105 µm (9.5 pixels per µm). However, this depends upon the Z-axis step size, the objective and the zoom factor.

A)	From the Stacks menu, select Options. Enter the appropriate value in the slice spacing dialog.

B)	Select an image from the series that best shows the features of interest. Using the Calibration/Measurement tool (dashed line in the right hand column of Tool Palette), draw a line across the section along the desired plane of resectioning. If you wish to constrain the plane to either the X-or Y-axis, press the Shift key and hold down as you draw the line.

C)	From the Stacks menu, choose Reslice (Command /). If circular profiles appear excessively flattened in either the X- or Y-axis, experiment with different values of "Slice Spacing" in the Options dialog.

E)	Reslicing To Make A Z-Series In An Alternate Plane

The preceding instructions describe how to obtain a single reslice in the X-Z or Y-Z plane. If you would like to obtain a new stack of Z-series section, not just a single section, use the macro provided, "Reslice Horizontally" or "Reslice Vertically". The reslicing is limited to orthogonal planes of section. Thus, if you wish to reslice along an oblique angle, you will first have to rotate the stack. In order to do this, use the Angle

Measurement tool from the Tool Palette. Measure the degrees of variance of the object in the image from the angle you finally desire.

1.	Use the macro "Scale and Rotate Stack", using the angle value established in the previous paragraph.

2.	Now select a rectangular area that encloses the region you want to reslice in either horizontal or vertical plane.

3.	Use the macro "Reslice Horizontally" or "Reslice Vertically".

1.	Rapid Dynamic 3D Reslicing

Norbert Vischer of the Netherlands has recently contributed an extremely useful macro, "3D Slicer" for reslicing a stack of sections along the X-, Y- and Z-axes. This macro is provided in the program "Object-Image 1.59", available from the NIH Image FTP site. Object-Image 1.59 provides extremely rapid reslicing in real time. The individual resliced sections cannot be saved, for unlike the macros described above, they are not placed in a new stack. The reslicing of sections is accomplished by dragging the mouse along the X-, Y- or Z-axes of a master stack, and results in rapid generation of resliced images in the two other planes. Reslicing can only be done along orthogonal planes.

A)	Open a Z-stack of sections, such as the sample MRI Images of a human skull and brain.

B)	Enter the correct magnification scale and slice spacing (5.0 mm), as described above.

C)	Under Stacks menu, select 3D Slicer. This will open a new window containing an angled perspective view of the original stack, with 3D-projection planes along the two alternate planes, as well. In the upper part of the window will be an image of the resliced section parallel to the X-axis. On the right side of the window will be the resliced section parallel to the Y-axis.

D)	Place the mouse over the X, Y or Z margin. The cursor changes to a two headed arrow indicating the direction of movement. Drag the X-, Y- or Z- axis on the main central image, and observe the rapid resectioning in the alternate planes.

E)	Place the mouse over the intersection of the X- and Y-axes and the cursor becomes a four-pointed arrow. You can now simultaneously reslice parallel to both the X- and Y-axes.

F)	Double clicking the mouse on the X-or Y-axes will turn off the reslicing tool for that axis. Double clicking on the dashed lines on the X- or Y-axis will restore the reslicing planes.

F)	Generating A 3D-Series From A Z-Series

This is a computationally intensive procedure. A fast PowerPC will prove very desirable when doing this operation. When you are still learning the basics of this procedure, use the Selection tool from the Tool Palette and outline a small portion of the image when generating a 3D-series. It will complete the task much more rapidly.

Use care in settings. Remember to correct for thickness and spacing of individual sections.

1.	Select Project from the Stacks menu.

Initially, follow these instructions:

2.	Set:

| Initial Angle (0-359): | 0 |
| --- | --- |

| Total Rotation (0-360): | 15 |
| --- | --- |

| Rotation Angle Increment: | 180 |
| --- | --- |

Leave the remaining values in their default configuration.

Select the "Y-Axis" as your "Axis of Rotation". This will cause the resulting image to rotate left to right (i.e., around  the Y-axis). Choose "Brightest Point" as the "Projection Method".

Now click OK and wait. The program will generate a new stack of images of the “rotating” objects in the original stack.

Once you have had further experience with this procedure, Project using a full 360 degree rotation at closer intervals, changing transparency values, axes of rotation and the various other options available.

A macro to facilitate this operation is provided.

G)	Animating A 3D-Series (Producing "Apparent Rotation")

Under Special menu, select Video Options. Check "Oscillate Movies". Close the dialog box.

To generate the impression of rotation with your newly generated 3D stack, use the Stacks menu and select Animate (Command+=). This is heavily memory dependent, so keep your initial efforts small.

You can control the speed of rotation with keys 1-9. You can step through single sections using the < and the > keys.

H)	Make A Stereo Pair Or Series Of Stereos?

There are several methods of generating stereo images from a stack of sections collected in the Z plane. One simple method, the "pixel shift" method, is that used in the original BioRAD software.

The pixel shift method starts with a stack of a Z-series. You will generate two Z-projections from this original stack. One will be "pixel shifted" to the left and a second "pixel shifted" to the right. This is achieved by shifting each successive section in the stack by one additional pixel (or more, if desired) prior to performing the Z-projection.  The two resultant images are then placed side by side. The pixel shift method is generally limited to stereo images centered around the original plane of image collection.

Projecting and ray-tracing is a second method that projects (ray-tracing) the stack onto an imaginary view plane from different angles of view and generating a 3D rotation

series. This second method is computationally more complex, but provides the possibility of generating stereo views from any angle around a central point.

Start with the original Z-series Stack. Select the Project from the Stacks menu and set "Initial Angle" to 354, "Increment  Angle" to 6 and "Total Rotation" to 12. This will produce a new Stack containing three slices (354°, 0° and 6°).

1.	Stereo Series In Black And White

Once you have generated a 3D-series, make a Montage series using the appropriate function from the Stacks menu. The Montage function is non-destructive, i.e., it does not erase the 3D stack, however, I suggest that you Save your work as you go along.

A)	Stacks Menu Montage

The resulting dialog box will show a series of values that are dependent upon the number of sections in the stack. If you made a simple stack with only 3 sections, then choose 1 row and 3 columns, and an increment of 1. If you generated a full 360° rotation series, you can select any range of slices (e.g., slices #4-8 out of 16 slices), or incremental slices (every 2nd or 3rd slice in the stack), and define the numbers of rows and columns you wish to see displayed.

This will generate a series of images in a new window on the screen. If the resulting images are too large to line up next to each other, use a scaling factor when generating the montage.

Start with two side-by-side images. To facilitate seeing the stereo effect without special viewers, choose a set of images with a prominent object to provide an alignment cue in the center of each image. Set the image size and separation with the alignment cues no more than 50 mm apart on the screen. Gradually work your way up to greater separations and then slightly larger images. Initially, you may find that you require a stereo viewer to visualize a stereo image from side-by-side images. With practice, you will not require a stereo viewer and should be able to scan across pairs of sections and jump from one stereo image to the next. (The typical interpupillary distance of most people is about 65 mm. You can also practice learning to fuse stereo images by using some of the recent popular books with "random dot stereograms", such as The Magic Eye.)

2.	Stereo Pair Of Single Labeled Section In Color

An alternate manner of presenting stereo images is to merge a stereo pair into a single image plane, assigning one image to red and the other to green.

To generate a 3D rotation series around the Y-axes., limit the number of image to two at a separation of ca. 6-12°. For the first trials, set the "Initial Angle" to 354, the "Increment Angle" to 6 and the "Total Rotation" to 12. This will generate three plates in a new stack. Animate the new stack of images, and if the animation produces the desired effect, delete either the middle or one of the end plates. Add a black plate as slice 3/3. (See section above on Merging Slices to generate 8-bit color images). Using the Stacks menu item RGB to 8-bit Color, you will obtain a single merged image which can be viewed with a pair of Red/Green stereo glasses.

The greater the separation between plates (i.e., 12° rather than 6°) the more exaggerated the stereo effect. Play with different values. Some people find that angles greater than 15-18° are excessive.

3.	Color Stereo Images Of A Double Labeled Section

Although this is a more complex set of operations, it is a logical extension of the methods described above.

First generate a 3D rotation series of each Z-series of a double labeled section. Save the results. Close all windows except those of the two 3D rotation series. Now use the macro "Color Merge of Two Stacks". You will now have a rotation series (3D) of two simultaneous different fluorophores. Animate the series.

Now generate stereo pairs, using the montage method described above.

I)	Exporting Stacks to QuickTime Movies and VCR Recordings of Stacks

A rotating Z-series, or any other Stack, can be saved in QuickTime format, allowing it to be viewed with a variety of other programs, such as a simple QuickTime Viewer, or with a program designed for preparing and transcribing clips on a VCR, such as Avid VideoShop 3.0. The latter program is presently distributed at no additional cost with many Apple Computers. Avid VideoShop can be used to prepare a VCR tape of your data, merging different data sets into a single sequence, etc.

1.	Saving Stacks In QuickTime Format

A)	Save the Stack in standard NIH Image format.

B)	Open the Stack, and select Save As...

C)	There are a series of optional buttons at the bottom of the dialog box, including TIFF, PICT, PICS and others.

D)	Choose "PICS"

E)	Rename file with new name (e.g., if file name was "NewCell", rename it "NewCell.PICS") to avoid overwriting your original data file.

F)	The new file can still be opened as a Stack within NIH Image.

G)	If you Save the file, it will revert to the original TIFF format, though with the new filename, and is likely to be confusing when you try to open it with a QuickTime player. Therefore, any additional Save operations should always be done using Save As....

An advantage of viewing the Stack with Avid VideoShop 3.0 is that you can increase the image size in continuous steps by dragging the lower right corner of the image window. This does not alter the aspect ratio of the original image.

VI.	Printing on Video Printer, Dye sublimation printer, Slide Maker And VCR

1.	Video Printer

You can obtain an instant print of the NIH Image window using a video printer, such as the Mitsubishi 67U or Sony thermal video printer. This provides a gray scale print. The Mitsubishi produces a print of approximately 3x4 inches, and provides a running record of your results. Scion markets a NuBus board (TV-3) that will directly send the contents of a selected window to an NTSC device. The driver for the TV-3 board must be placed in the NIH Image Plug-Ins folder. The drawback of the TV-3 is that it can only send out an image of maximum dimension of 640x480 pixels. Since the standard BioRAD image is 768x512, the image must first be scaled to a 640x480 size, or select a portion of the image corresponding to this size. Images smaller than 640x480 will only fill a portion of the printed field. It would be useful if the driver provided an autoscaling function. A simple macro, however, can be written to accomplish a similar result. The new PowerPC 8500 provides direct NTSC and S-Video output. I have not had a chance to test them with the Mitsubishi printer, but based on published specifications, these should provide results similar to those obtained with the Scion TV-3 board.

2.	Dye Sublimation Printer

Prints of much better quality, including black and white or color prints, can be obtained using the Kodak 8600 dye sublimation printer. The printer is provided with a Adobe Photoshop Plug-In driver. Place this Plug-In in the folder of the same name in the main NIH Image directory. When you want to print the contents of a window, select the desired image. Under the File menu, choose Export, and select the "Kodak 8600" printer driver. The typical image will print in approximately one minute. If you are in NIH Image, this works only with monochrome images printed on a black ribbon. Color prints are best produced using Adobe Photoshop. A major limitation in the use of the Kodak printer is the cost of each print ($2-$3), as well as the cost of the printer itself (ca. $8,000).

The wider availability of color laser printers may provide satisfactory color prints at lower cost than those obtained with the Kodak 8600 printer.

3.	Slide Maker

There are a number of slide makers, including the LFR Lasergraphics, the GCC/Polaroid slidemaker, Agfa and others.

If you are using Microsoft PowerPoint to send images to a slidemaker, Scale the images to the full size of the screen (e.g., 768x512) prior to saving them. Use the "Bilinear Interpolation", rather than "Nearest Neighbor", for best results. You can also use PowerPoint to combine different images on a single slide, add text, or other graphics elements. Paste the NIH Images onto a black background to avoid having a white border around the image.

4.	VCR

Videotape players are now commonly available at all universities and scientific meetings. You may find it far more effective to display rotating projections as well as sequences of sections on videotape, rather than with 2x2 slides or overheads. The technology for transferring such Stacks is widely available, and a brief description of its application to NIH Image and confocal Stacks is described in the preceding section.

VII.	USE OF NIH IMAGE FOR IMAGE COLLECTION AND INSTRUMENT CONTROL

The majority of CLSMs have their own dedicated software for image collection. However, several instruments, such as BioRAD's direct viewing slit scanner, and Noran's AOD real-time scanner have also been used directly with NIH Image for data collection.

VIII.	APPENDIX

1.	Macros

A set of macros has been provided for use with confocal images, as well as for the collection of fluorescent images using an integrating CCD. These are to be found in a separate file, entitled "Image Macros for CLSM". They can be used directly with NIH Image without modification. However, you may find that your personal preferences differ from mine, and you may modify the macros to better suit your needs. See Mark Vivino's excellent manual on Macro Programming with NIH Image, available from the usual FTP site (zippy.nimh.nih.gov).

If you wish these macros to load automatically each time you start NIH Image, place them in the same folder with the program or in the System Folder. Before you do that, Save the current version of "Image Macros" using an alternate name. Rename a copy of  "Image Macros for CLSM" as "Image Macros".

The next time you start NIH Image, these macros will load as the default set.  See the NIH Image manual for further details on loading macros.

IX.	ADDENDA TO NIH IMAGE CONFOCAL IMAGING MANUAL

Addenda March, 1995

1.	Macros to merge sides of split screen BioRAD image. Macro to swap red and green slices of an RGB Stack.

2.	Macro to color merge Z-series of split screen BioRAD images.

3.	Notes on exporting/importing Adobe Photoshop format.

4.	Comments on scaling using "Bilinear" rather than "Nearest Neighbor".

Addenda June, 1995

1.	Transferring NIH Image RGB images to Adobe Photoshop 3.0

2.	New macro to automate Z-projection

Addenda September, 1996

1.	Use of NIH Image for pre-evaluation of fluorescence--shutter and on-chip integration control. Provides a series of macros for controlling the shutter, on-chip integration, making new Stacks and related operations.

2.	Database storage of images--put Z-projection image on top of Stack for databasing. Compare FileMaker Pro, Search 3.1, Cumulus 2.5, and Kudo.

3.	GetPath functions--a means of imprinting the file name and full path on each image is provided in NIH Image version 1.60. This greatly simplifies record keeping.

4.	Notes on Z-projections

5.	Caution on rotation

6.	Revision on recommendation for storage media.

List of Macros provided in "Confocal Macros"

These macros are arranged in order of my personal preference (more or less). This order will probably not satisfy anyone else. If you rearrange the order of these macros, make sure that you are careful in regard to defining "Variables and Procedures". They must all be placed in the file in front of the macros that call them.

Define Procedures for Stacks

procedure Error(s:string);

procedure CalibrateImage

procedure CheckForStack;

procedure CheckForSelection;

procedure GetSliceRange(smart:boolean);

procedure GetSliceSequence(smart:boolean);

Some Universal Operations

macro '[P] Print Video';

macro '[Q] Get Path';

macro '[F2] Calibrate Image';

Operations on Confocal Files: Import, Split, Merge

macro [=] "Animate";

procedure ShowBioRadInfo(InfoOffset: integer);

macro '[F1] Import Biorad Z Series';

macro '[S] Merge Split BioRAD';

macro ' Color Merge Two Images';

macro ' Color Merge Two Stacks';

macro 'Separate SplitScreen Z Stack';

macro '[F5] RGB to Indexed ';

macro '[W] Swap Red_Green';

macro ' Merge Two Stacks';

General Stack Operations

macro ' Make stack size front image';

macro '[F8] Flash to Stack

macro MakeStack w_Current Image';

macro 'Add Slice';

macro '[A] Add Black Slice';

macro '[D] Delete Slice';

Shutter Controls and On-Chip Integration

macro ' Autoshutter';

procedure ResetShutter;

macro '[O] Open Shutter';

macro '[C] Close Shutter';

macro '[T] Trigger Shutter';

macro '[R] Reset Shutter';

macro '[G] Open Shutter-Grab-Close';

procedure EndIntegration;

procedure Integrate (mode:string);

macro '[F6] Integrate On-chip Using Cohu';

macro '[F7] Integrate One Image on Cohu';

macro '[2] SetIntegrate:2 Frames';

macro '[F] Show nFrames';

macro 'Integrate White Light';

macro ' Live';

macro ' Average';

macro '[F1] Import Biorad Z Series';

macro '[F2] Calibrate Image';

Stack Modifications, Z-Projections, Stereo Pairs, etc.

macro ' Smooth Stack';

macro ' Sharpen Stack';

macro '[I] Invert Stack';

macro ' Reduce Noise Stack';

macro '[L] Apply LUT to Stack ';

macro '[F0] Remove 0 and 255 from Stack';

macro ' Flip Stack Vertical';

macro ' Flip Stack Horizontal';

macro ' Clear Outside Stack';

procedure CropAndScale(fast:boolean; angle:real);

macro '[E] Crop and Scale-Fast…';

macro ' Crop and Scale-Smooth…';

macro ' Rotate Left';

macro ' Rotate Right';

macro ' Rotate…';

procedure DoReslicing(horizontal:boolean);

macro '[H] Reslice Horizontally';

macro '[V] Reslice Vertically';

macro '[F12] Z Projection';

macro '[F14] Make Stereo Views';

LUT Manipulations

macro ' Invert LUT';

macro ' Log Tranform';

macro '[M] Gamma Tranform…';

macro ' Reset LUT';

macro ' Set Pixels Red…';

macro ' Nearly Gray LUT…';

macro '[U] Move Slice Up';

macro '[Y] Move Slice Down ';

Hot Keys

0 - macro '[0] Remove 0 and 255 from Stack';

2 - macro '[2] SetIntegrate:2 Frames';

6 - macro 'Make 640x480 ROI';

= - macro 'Animate';

A - macro '[A] Add Black Slice';

C - macro '[C] Close Shutter';

D - macro '[D] Delete Slice';

E - macro '[E] Crop and Scale Stack-Fast';

F - macro '[F] Show nFrames';

G - macro '[G] Open Shutter-Grab-Close';

H - macro '[H] Reslice Horizontally';

I - macro '[I] Invert Stack';

L - macro '[L] Apply LUT to Stack ';

M - macro '[M] Get Rectangle...';

N - macro '[N] Make Stack Same Size as Front Image';

O - macro '[O] Open Shutter';

P - macro '[P] Print Video';

Q - macro '[Q] Get Path';

R - macro '[R] Reset Shutter';

S - macro '[S] Merge Split BioRAD';

T - macro '[T] Trigger Shutter';

U - macro '[U] Move Slice Up';

V - macro '[V] Reslice Vertically';

W - macro '[W] Swap Red_Green;

X - macro '[X] Gamma Tranform...';

Y - macro '[Y] Move Slice Down ';

F1- macro '[F1] Import Biorad Z Series';

F2 - macro '[F2] Calibrate Image';

F3 - macro '[F3] Z Projection';

F4 - macro'[F4] Make Stereo View';

F5 - macro '[F5] RGB to Indexed '

F6 - macro '[F6] Integrate On-chip Using Cohu';

F7 - macro '[F7] Integrate One Image on Cohu';

F8 - macro '[F8] Flash to Stack';

F10 macro '[F10]   Choose First Slice';

F11 macro '[F11]   Choose Last Slice';

F12 - macro '[F12] Z Projection of Partial Stack';

Macros

{The following Macros should be copied to a separate file and saved in a simple text format, in order to be accessible to NIH Image. }

{Copy all the text from this point to the end of the file into a text file. }

{The resultant file should be saved as "Confocal Macros"}

{If you want these macros to load automatically when you start NIH Image,}

{rename them as "Image Macros" and place them in the same directory }

{that contains the program NIH Image.}

{BEGIN Macros HERE}

var

top,left,n,w,h,width,height:integer;

name: string;

{ Global variables }

var

n,nFrames:integer; {Global variable used by integration macros}

mag, barlength: real;	{ objective magnification, size of scale bar to use (µm) }

camerapid: integer;	{ PID of camera window }

bgpid: integer;		{ background window }

bgsub: boolean;		{ whether to subtract background }

tsoff: boolean;		{ whether to omit timestamp from captured images }

asoff: boolean;		{ whether the AutoShutter feature is off }

shopen: boolean;		{ whether we think the shutter is open }

tlperiod: real;		{ timelapse period }

name: string;

StackName:string;

OriginalStack,FirstSlice:Integer;

LastSlice,NextSlice,PartialStack:Integer;

left,top,x1,y1,x2,y2,lw:Integer;

{ Error(s) issues the error message S and terminates the macro. }

procedure Error(s:string);

begin

PutMessage(s); exit;

end;

procedure CheckForStack;

begin

if nSlices=0 then begin

PutMessage('This window is not a stack');

exit;

end;

end;

procedure CheckForSelection;

var x1,y1,x2,y2,lw,l,t,w,h:integer;

begin

GetRoi(l,t,w,h);

GetLine(x1,y1,x2,y2,lw);

if (w=0) and (x1<0) then Error('Please make a rectangular selection.');

end;

procedure ShowRange;

{Display file name, PidNumber, number of slices in stack, first and last slices to be used.}

begin

ShowMessage('FileName: ' StackName ,'\ Image Size = 'width ' x ' height ' pixels' ';\ Number of Slices = ',nSlices, ';\ First Slice # = ', FirstSlice,';\ Last Slice # = ' LastSlice);

end;

end;

{

GetSliceRange is a procedure that gets the implied parameters first and last,

then does range checking to make sure this is a valid range of slices.

GetSliceSequence gets the implied parameters first, last, and step, doing range checking to make sure it's a valid sequence of slices.

}

procedure GetSliceRange(smart:boolean);

var default:integer;

begin

first:=GetNumber('First slice:',1);

if (first<1) or (first>nSlices) then Error('Not a valid slice number.');

default:=nSlices;

if smart and (first=1) and (SliceNumber > 1) then default:=SliceNumber;

last:=GetNumber('Last slice:',default);

if (last<first) or (last>nSlices) then Error('Not a valid range of slices.');

end;

procedure GetSliceSequence(smart:boolean);

begin

GetSliceRange(smart);

step:=GetNumber('Step:', 1);

if (step<1) then Error('Step must be 1 or greater.');

end;

procedure CalibrateImage;

var

unit: string;

SliceSpacing,PixPerMicron,ActualSpacing: real

begin

SetScale(0,'');

GetScale(PixPerMicron,unit);

if nSlices>1 then begin

SliceSpacing :=GetNumber('Section Interval:',0.35,2);

ActualSpacing:= (SliceSpacing*PixPerMicron);

SetSliceSpacing(ActualSpacing);

exit; end;

end;

procedure Animate;

{Not as fast as the Animate command under the Stacks Menu (Command+=)}

var

i,delay:integer;

begin

CheckForStack;

RequiresVersion (1.50);

i:=0;

delay:=0.1;

repeat

i:=i+1;

if i>nSlices then i:=1;

Wait(delay);

SelectSlice(i);

if KeyDown('shift') then delay:=1.5*delay;

if delay>1 then delay:=1;

if KeyDown('control') then delay:=0.66*delay;

if KeyDown('option') then exit;

ShowMessage('delay=', delay:4:6);

until button;

end;

macro '[M]      Get Rectangle'; begin SelectTool('rectangle'); end;

macro '[6]       640x480 ROI';

begin

MakeRoi(0,0,640,480);

end;

macro '[P]       Print Video';

begin

CallExport('TV-3 Module');

end;

macro '[Q]       Get Path';

var

wPath, sPath, pPath: string;

name, FullPath, FileType, folder: string;

FileSize: integer;

begin

name := WindowTitle; {Grab name before before opening window}

wPath := GetPath('window'); {Grab path before before opening window}

SaveState;

SetFont('Geneva');

SetFontSize(12);

NewTextWindow('Path:',550,75);

if wPath = '' then begin

writeln('Image not saved with unique FileName, ');

writeln(' or No opened image or text window');

end else begin

FullPath := concat(wPath, name);

{writeln('Active window path = ');}

writeln( FullPath);

SelectAll;

Copy;

end;

end;

RestoreState;

end;

macro '(-';begin end;

procedure ShowBioRadInfo(InfoOffset: integer);

{Displays the contents of the  'header' at}

{the end of Biorad MRC 600/1000 Lasersharp 1024  single section files.}

{Needs correction of Offset values to properly read Z-series info}

var

MaxInfoSize,offset:integer;

ch, title:string;

begin

MaxInfoSize:=4096;

SetCustom(MaxInfoSize, 1, InfoOffset);

SetImport('8-bit'); {Don't invert}

Import('');

GetRow(0,0,MaxInfoSize);

Dispose;

SaveState;

title := concat(WindowTitle, '.Info');

NewTextWindow(title, 500,400);

SetCursor('Watch');

SetFont('Monaco');

SetFontSize(12);

for i:=0 to MaxInfoSize-1 do begin

offset:=i mod 96;

if offset=0 then writeln;

ch:=chr(LineBuffer[i]);

if (offset=2) and (ord(ch)=0) then exit;

if (offset>=16) and (offset<=95) and (ord(ch)>=32) and (ord(ch)<=126)

then write(ch);

end;

RestoreState;

end;

macro '[F1]      Import Biorad Z Series';

{macro 'Import Biorad  Z Series…';}

{

Imports a Z series(multiple images per file) from a Biorad MRC 600/1000

confocal microscope.  The width, height and number of images are

extracted from the first 3 16-bit word in the 76 byte header and

the file name is extracted from bytes 18-23 of the header. This macro

does not read merged pseudocolored BioRad files. Note that the Undo

and Clipboard buffers must be set to 384K to work with the typical

768x512 Biorad images.

}

var

width,height,nImages,offset,hdrsize,i,start,picsize:integer;

begin

RequiresVersion(1.50);

width:=512;

height:=1;

offset:=0;

SetImport('8-bit');

SetCustom(width,height,offset);

Import(''); {Read header}

GetPicSize(width,height);

if (width<>512) or (height<>1) then begin

Dispose;

PutMessage('Please to not change width, height, etc. in the Import dialog box.');

exit;

end;

width:=GetPixel(0,0)+GetPixel(1,0)*256;

height:=GetPixel(2,0)+GetPixel(3,0)*256;

nImages:=GetPixel(4,0)+GetPixel(5,0)*256;

Dispose;

hdrsize:= 76;

picsize:=width*height;

if (width<128) or (width>2048) or (height<128) or (height>2048) or (nImages<1) or (nImages>256) then begin

PutMessage('This does not seem to be a Biorad MRC 600/1000 Z Series file.');

exit;

end;

start:=GetNumber('Starting image:',1);

offset:=HdrSize+(start-1)*PicSize;

SetImport('8-bit, Invert');

SetCustom(width,height,offset,nimages);

Import('');

ShowBioRadInfo(HdrSize + nImages * width * height);

CalibrateImage;

end;

macro '[F2]     Calibrate Image';

begin CalibrateImage; end;

macro '[F3]      Z Projection';

Begin

SetProjection('Initial Angle',0);

SetProjection('Total Rotation',0);

SetProjection('Rotation Increment', 0);

SetProjection('Interior Depth-Cueing', 0);

SetProjection('Y-Axis');

SetProjection('Brightest');

SetProjection('Save Projections', false);

SetProjection('Minimize Size', false);

SetDensitySlice(0,254);

Project;

end;

macro'[F4]      Make Stereo Views';

Begin

SetProjection('Initial Angle',348);

SetProjection('Total Rotation',24);

SetProjection('Rotation Increment', 6);

SetProjection('Interior Depth-Cueing', 0);

SetProjection('Y-Axis');

SetProjection('Brightest');

SetProjection('Save Projections', false);

SetProjection('Minimize Size', false);

SetDensitySlice(0,254);

Project;

Animate;

end;

macro '[F5]     RGB to Indexed ';

begin

RGBToIndexed('System, Dither');

end;

macro '[W]      Swap Red_Green';

begin

CheckforStack;

ChooseSlice(1);

SelectAll;

Copy;

DeleteSlice;

SetBackground(255);

AddSlice;

Paste;

KillRoi;

RGBToIndexed('System, Dither');

end;

macro '[S]       Merge Split BioRAD';

var

OriginalPic,w,h,rgb: integer;

{name: string;}

begin

OriginalPic:= PidNumber;

SelectPic(OriginalPic);

GetPicSize(w,h);

MakeRoi(0,0,w/2,h);

Copy;

SetNewSize(w/2,h);

name := GetString('New Stack Name', 'RGB');

MakeNewStack(name);

rgb:=PicNumber;

SelectPic(rgb);

Paste;

SetBackground(255);

AddSlice;

AddSlice;

SelectPic(OriginalPic);

MakeRoi(w/2,0,w/2,h);

Copy;

KillRoi;

SelectPic(rgb);

SelectSlice(2);

Paste;

SetBackGround (255);

RGBToIndexed('System,Dither');

end;

macro '           Color Merge Two Images';

var

i,w1,w2,h1,h2,rgb:integer;

begin

RequiresVersion(1.50);

SaveState;

if nPics<>2 then begin

PutMessage('This macro operates on exactly two images.');

exit;

end;

SelectPic(1);

GetPicSize(w1,h1);

SelectPic(2);

GetPicSize(w2,h2);

if (w1<>w2) or (h1<>h2) then begin

PutMessage('The two images must have the same width and height.');

exit;

end;

SetNewSize(w1,h2);

SetBackground(255);

MakeNewStack('RGB');

AddSlice;

AddSlice;

rgb:=PicNumber;

SelectPic(1);

SelectAll;

Copy;

SelectPic(rgb);

SelectSlice(1);

Paste;

SelectPic(2);

SelectAll;

Copy;

SelectPic(rgb);

SelectSlice(2);

Paste;

RGBToIndexed('Custom');

end;

macro '           Color Merge Two Stacks';

var

i,w1,w2,h1,h2,d1,d2,d3:integer;

rgb,merged:integer;

begin

RequiresVersion(1.50);

SaveState;

if nPics<>2 then begin

PutMessage('This macro operates on exactly two stacks.');

exit;

end;

ChoosePic(1);

GetPicSize(w1,h1);

d1:=nSlices;

ChoosePic(2);

GetPicSize(w2,h2);

d2:=nSlices;

if (d1=0) or (d2=0) then begin

PutMessage('Both images must be stacks.');

exit;

end;

if d1>=d2

then d3:=d2

else d3:=d1;

if (w1<>w2) or (h1<>h2) then begin

PutMessage('The two stacks must have the same width and height.');

exit;

end;

SetNewSize(w1,h2);

SetBackground(255);

MakeNewStack('RGB');

AddSlice;

AddSlice;

rgb:=PicNumber;

SetPalette('System');

MakeNewStack('Merged');

merged:=PicNumber;

for i:=1 to d3 do begin

ChoosePic(1);

ChooseSlice(i);

SelectAll;

Copy;

{ Following line was deleted,  as it  makes value of 'd3' erroneous when used in loop}

{DeleteSlice;}

ChoosePic(rgb);

ChooseSlice(1);

SelectAll;

Paste;

{ Invert;}

ChoosePic(2);

ChooseSlice(i);

SelectAll;

Copy;

{Following line was deleted, for same as previous reason above }

{ DeleteSlice;}

ChoosePic(rgb);

ChooseSlice(2);

SelectAll;

Paste;

{ Invert;}

SelectPic(rgb);

RGBToIndexed('System,Dither');

SelectAll;

Copy;

Dispose;

ChoosePic(merged);

Paste;

if i<>d3 then AddSlice;

end;

ChoosePic(rgb);

Dispose;

RestoreState;

end;

macro '           Separate SplitScreen Z Stack';

var

LeftStack,RightStack,OriginalStack,w,h,i,OriginalnSlices:integer

begin

{set up parameters for new stacks}

CheckForStack;

OriginalStack:= PidNumber;

ChoosePic(OriginalStack);

OriginalnSlices:=nSlices;

GetPicSize(w,h);

SetNewSize(w/2,h);

SetBackground(255);

name := GetString('New Stack Name', 'RGB');

MakeNewStack(name,' Left Stack');

LeftStack:=PidNumber;

MakeNewStack(name,' Right Stack');

RightStack:=PidNumber;

{ OK, now you have two stacks, Left Stack and Right Stack}

{ PutMessage('This Stack has ',nSlices,' slices. Slice =',i);}

for i:= 1 to OriginalnSlices do begin

ChoosePic(OriginalStack);

SelectSlice(i);

MakeRoi(0,0,w/2,h);

Copy;

ChoosePic(LeftStack);

ChooseSlice(i);

Paste;

SetBackground(255);

AddSlice;

ChoosePic(OriginalStack);

MakeRoi(w/2,0,w/2,h);

Copy;

KillRoi;

ChoosePic(RightStack);

SelectSlice(i);

Paste;

SetBackground(255);

AddSlice;

KillRoi;

{  SelectPic(OriginalStack);}

end;

end;

macro '           Merge Two Stacks';

{

Combines two stacks(w1xh1xd1 and w2xh2xd2) to create a new

w1+w2 x max(h1,h2) x max(d1,d2) stack. For example, a 256x256x40

and a 256x256x30 stack would be combined into one 512x256x40 stack.

}

var

i,w1,w2,w3,h1,h2,h3,d1,d2,d3:integer;

begin

SaveState;

if nPics<>2 then begin

PutMessage('This macro operates on exactly two stacks.');

exit;

end;

SelectPic(1);

GetPicSize(w1,h1);

d1:=nSlices;

SelectPic(2);

GetPicSize(w2,h2);

d2:=nSlices;

if d1>=d2

then d3:=d1

else d3:=d2;

if d3=0 then begin

PutMessage('Both images must be stacks.');

exit;

end;

w3:=w1+w2;

if h1>=h2

then h3:=h1

else h3:=h2;

SetNewSize(w3,h3);

MakeNewStack('Merged');

for i:=1 to d3 do begin

SelectPic(1);

SelectSlice(1);

SelectAll;

Copy;

DeleteSlice;

SelectPic(3);

MakeRoi(0,0,w1,h1);

Paste;

SelectPic(2);

SelectSlice(1);

SelectAll;

Copy;

DeleteSlice;

SelectPic(3);

MakeRoi(w1,0,w2,h2);

Paste;

if i<d3 then AddSlice;

end;

SelectPic(1);

Dispose;

SelectPic(1);

Dispose;

RestoreState;

end;

macro '(-' begin end;

macro '[F10]   Choose First Slice';

{1) Choose the Stack you want to fractionate.

2) grab the WindowTitle, PidNumber, Window or ROI size in pixels.

3) select the beginning slice - using < or > keys and hit key 'F'. This will  store  slicenumber as  variable FirstSlice.

Easily modified to inform  user of  size of  resulting Stack in bytes.

Proceed macro Choose LastSlice.}

}

begin

CheckForStack;

StackName:=WindowTitle;

OriginalStack:=PidNumber;

GetRoi(left,top,width,height);

if (width=0)  then

begin

SelectAll;

GetPicSize(width,height);

KillRoi;

end;

{CheckForSelection;}{Replace following SelectAll with ROI?}

FirstSlice:=SliceNumber;

ShowRange;

{SelectAll;}

end;

macro '[F11]   Choose Last Slice';

{ Select the last slice that is to be used to end the new stack

and hit key 'L'}

begin

CheckforStack;

OriginalStack:=PidNumber;

StackName:=WindowTitle;

LastSlice:=SliceNumber;

ShowRange;

end;

macro '              Show range of slices';

begin

ShowMessage('"' StackName'"' ,' has PidNumber = ', OriginalStack,'\ Image Size = 'width ' x ' height ' pixels' ';\ Number of Slices = ',nSlices, ';\ First Slice # = ', FirstSlice,';\ Last Slice # = ' LastSlice);

If FirstSlice>LastSlice then

PutMessage('Selection of First Slice must precede Last Slice');

end;

end;

procedure CheckSliceRange;

Begin

If FirstSlice>LastSlice then begin

PutMessage('Not a valid range of slices. First Slice must precede Last Slice in Stack');

exit;

end;

end;

macro '[F12]   Z-Projection of Partial Stack';

begin

CheckSliceRange;

SetNewSize(width,height);

SetBackground(255);

MakeNewStack('Partial Stack');

PartialStack:=PidNumber;

Begin

SelectPic(OriginalStack);

For NextSlice := FirstSlice to LastSlice do		begin

SelectSlice(NextSlice);

MakeROI(left,top,width,height);

Copy;

SelectPic(PartialStack);

Paste;

AddSlice;

SelectPic(OriginalStack);

NextSlice:=NextSlice+1;

end;

Killroi;

end;

Begin

SelectPic(PartialStack);

CheckForStack;

SetProjection('Initial Angle',0);

SetProjection('Total Rotation',0);

SetProjection('Rotation Increment', 0);

SetProjection('Interior Depth-Cueing', 0);

SetProjection('Y-Axis');

SetProjection('Brightest');

SetProjection('Save Projections', false);

SetProjection('Minimize Size', false);

SetDensitySlice(0,254);

Project;

end;

end;

macro '(-';

{macro 'Make stack same size as front image                 [N]';

var

width, height: integer;

name: string;

begin

SaveState;

GetPicSize(width, height);

SelectAll;

Copy;

KillRoi;

SetNewSize(width, height);

name := GetString('New Stack Name', 'stack');

SetBackground(255);

MakeNewStack(name);

Paste;

KillRoi;

RestoreState;

end;}

macro '           Make stack size front image';

var

width, height: integer;

begin

SaveState;

GetPicSize(width, height);

SetNewSize(width, height);

name := GetString('New Stack Name', 'stack');

MakeNewStack(name);

RestoreState;

end;

macro '           MakeStack w_Current Image';

var

w,h:integer;

begin

SelectAll;

GetPicSize(w,h);

SetNewSize(w,h);

Copy;

MakeNewStack('NewStack');

Paste;

KillRoi;

SetBackground(255);

AddSlice;

AddSlice;

SelectSlice(2);

end;

macro '           Add Slice';    begin AddSlice end;

macro '[A]       Add Black Slice';

begin

SetBackground(255);

AddSlice;

end;

macro '[D]       Delete Slice'; begin DeleteSlice end;

macro '[=]       Animate Stack';

{Hit Shift to slow down, Control to speed up and Option to halt}

begin animate;end;

macro '(-' begin; end;

macro '            Autoshutter';

begin

asoff := not asoff;

if asoff then ShowMessage('Autoshutter off.')

else ShowMessage('Autoshutter on.');

end;

procedure CloseShutter;

begin

OpenSerial('9600 baud,no parity,eight data,one stop');

PutSerial(chr(49));

shopen:=false;

end;

procedure OpenShutter;

begin

OpenSerial('9600 baud,no parity,eight data,one stop');

PutSerial(chr(48));

shopen:=true;

end;

procedure TriggerShutter;

begin

OpenSerial('9600 baud,no parity,eight data,one stop');

PutSerial(chr(50));

shopen:=false;

end;

procedure ResetShutter;

begin

OpenSerial('9600 baud,no parity,eight data,one stop');

PutSerial(chr(51));

shopen:=false;

end;

macro '[O]        Open Shutter'; begin OpenShutter; end;

macro '[C]        Close Shutter'; begin CloseShutter; end;

macro '[T]        Trigger Shutter'; begin TriggerShutter; end;

macro '[R]        Reset Shutter'; begin ResetShutter; end;

macro '[G]        Open Shutter-Grab-Close';

begin

OpenShutter;

StartCapturing;

CloseShutter;

end;

procedure EndIntegration;

begin

CloseShutter;

Exit;

end;

procedure Integrate (mode:string);

var

x,y,delta:integer;

begin

if nFrames=0 then nFrames:=2;

SelectWindow('Camera');

repeat

if button then begin

GetMouse(x,y);

if (x<0) or (y<0) then EndIntegration;

delta:=round(0.250*nFrames);

if delta<1 then delta:=1;

if y>220 then begin

nFrames:=nFrames-delta;

if nFrames<1 then nFrames:=1;

end else begin

nFrames:=nFrames+delta;

if nFrames>240 then nFrames:=240;

ShowHistogram;

end;

SelectWindow('Camera');

end;

AverageFrames(mode, nFrames);

ShowMessage('# Frames Integrated = ',nFrames);

until false;

end;

macro '[F6]      Integrate On-chip Using Cohu';

begin

OpenShutter;

Integrate('integrate on-chip'); CloseShutter;

end;

macro '[F7]      Integrate One Image on Cohu';

begin

SelectWindow('Camera');

OpenShutter;{Wait(0.1);}

AverageFrames('integrate on-chip',nFrames);

CloseShutter;

{Add code to for TimeStamp on window}

{ShowHistogram;}

ShowMessage('# Frames Integrated = ',nFrames);

end;

{Add macro to Re-Set number of frames to 2}

macro '[2]        SetIntegrate:2 Frames';

begin

nFrames:=2;

end;

macro '[F]        Show nFrames';

begin

{PutMessage('# Frames Integrated = ',nFrames);}

ShowMessage('# Frames Integrated = ',nFrames);

end;

macro '             Integrate White Light';

begin

Integrate('integrate on-chip');

end;

{ Live and Average are fancy context-sensitive macros:

• in the Camera window, Live starts capturing and Average adds to a stack

• in another window with a ROI selected, Live starts a Live paste, and Average

pastes in an averaged picture for the ROI

• in another window with no ROI, Live switches to the Camera window, and starts

capturing

}

macro '             Live';

var l,t,w,h:integer;

begin

GetROI(l,t,w,h);

RequiresVersion(1.53);

if not asoff then OpenShutter;

if (camerapid=0) or (WindowTitle='Camera') or (w=0) then begin

StartCapturing;

camerapid := PidNumber; exit;

end;

{ If we get to here, we're in a non-Camera window with a ROI; start Live Paste}

PasteLive;

end;

macro '             Average';		{ assumes that Live has been called already }

var

l,t,w,h,destpid,p:integer;

year,month,day,hour,minute,second,DoW:integer;

camera:boolean;

begin

GetROI(l,t,w,h);

RequiresVersion(1.53);

if not asoff and not shopen then begin

OpenShutter; Wait(.1);		{ Compensate for SIT's lag time }

end;

destpid:=PidNumber;

camera:=(WindowTitle='Camera');

if camera then begin	{ try to find the first stack }

p:=0;

repeat

p:=p+1;

Choosepic(p);

until (nSlices>0) or (p>=nPics);

if (nSlices>0) then destpid:=p;

{ else leave destpid pointing at Camera, later code will start a new stack }

Choosepic(camerapid);

end;

if not camera then begin

ChoosePic(camerapid);

if (w>0) then MakeROI(l,t,w,h);

end;

AverageFrames;

if (w=0) then SelectAll;

Copy;

GetTime(year,month,day,hour,minute,second,DoW);

if not asoff then CloseShutter;

SelectPic(destpid);

if (w>0) and not camera then begin		{ we were in PasteLive mode }

MakeRoi(l,t,w,h); Paste; BkgdSub; exit;

end;

if nSlices>0 then begin

SelectSlice(nSlices);

AddSlice;

end;

if nSlices=0 then begin

if (camera and (w>0)) then SetNewSize(w,h) else SetNewSize(640,480);

MakeNewStack(month:2,'/',day:2,'/',year-1900:2,' ',hour:2,'.',minute:2,'.',second:2);

end;

Paste; BkgdSub;LabelFrame(0,0,not tsoff);

if (camera and (w>0)) then SelectPic(camerapid);

end;

{

Flash to Stack flashes once and puts the averaged picture in the first stack

that it finds. If there's no stack, it creates one. If you frame a ROI in the

camera window, only the ROI contents will be saved in the stack...

}

macro '[F8]       Flash to Stack ';

var

l,t,w,h,destpid,p:integer;

year,month,day,hour,minute,second,DoW:integer;

camera:boolean;

begin

GetROI(l,t,w,h);

RequiresVersion(1.53);

destpid:=PidNumber;

camera:=(WindowTitle='Camera');

if (camerapid=0) then begin

Capture; camerapid := PidNumber;

if (destpid=0) then destpid:=camerapid else ChoosePic(destpid);

end;

if camera then begin	{ try to find the first stack }

p:=0;

repeat

p:=p+1;

ChoosePic(p);

until (nSlices>0) or (p>=nPics);

if (nSlices>0) then destpid:=p;

{ else leave destpid pointing at Camera, later code will start a new stack }

Choosepic(camerapid);

end;

if not camera then begin

ChoosePic(camerapid);

if (w>0) then MakeROI(l,t,w,h) else SelectAll;

end;

OpenSerial('9600 baud'); StopCapturing; { in case we're live }

{ wait for even frame of the video signal, then flash and grab }

repeat until ((BitAnd(Scion[3],48)=0));

PutSerial(chr(48));

{ AverageFrames('Video Rate Capture',2); }

{test integration with shutter}

Integrate('integrate on-chip');

CloseShutter;

GetTime(year,month,day,hour,minute,second,DoW);

if (w=0) then SelectAll; Copy;

SelectPic(destpid);

if nSlices>0 then begin

SelectSlice(nSlices);

AddSlice;

end;

if nSlices=0 then begin

if (camera and (w>0)) then SetNewSize(w,h) else SetNewSize(640,480);

MakeNewStack(month:2,'/',day:2,'/',year-1900:2,' ',hour:2,'.',minute:2,'.',second:2);

end;

Paste; BkgdSub; LabelFrame(0,0,not tsoff);

if (camera and (w>0)) then SelectPic(camerapid);

end;

macro '(-'; begin; end;

macro '              Smooth Stack';

var

i:integer;

begin

CheckForStack;

for i:= 1 to nSlices do begin

SelectSlice(i);

SetOption; Smooth;

end;

end;

macro '              Sharpen Stack';

var

i:integer;

begin

CheckForStack;

for i:= 1 to nSlices do begin

SelectSlice(i);

SetOption; Smooth;

SetOption; Sharpen;

end;

end;

macro '[I]          Invert Stack';

var

i:integer;

begin

CheckForStack;

for i:= 1 to nSlices do begin

SelectSlice(i);

Invert;

end;

end;

macro '              Reduce Noise Stack';

var

i:integer;

begin

CheckForStack;

for i:= 1 to nSlices do begin

SelectSlice(i);

ReduceNoise;

end;

end;

macro '[L]         Apply LUT to Stack ';

var

i,stack,slices:integer;

begin

CheckForStack;

stack:=PicNumber;

slices:=nSlices;

Duplicate('Temp');

for i:= 1 to slices do begin

SelectPic(stack);

SelectSlice(i);

ApplyLut;

SelectPic(nPics);

if i<>slices then PropagateLut;

end;

SelectPic(nPics);

Dispose;

end;

macro '[0]         Remove 0 and 255 from Stack';

var

i:integer;

begin

CheckForStack;

for i:= 1 to nSlices do begin

SelectSlice(i);

ChangeValues(0,0,1);

ChangeValues(255,255,254);

end;

end;

procedure flip(vertical:boolean);

var

i:integer;

begin

CheckForStack;

for i:= 1 to nSlices do begin

SelectSlice(i);

if vertical

then FlipVertical

else FlipHorizontal;

end;

end;

macro '              Flip Stack Vertical';   begin flip(true) end;

macro '              Flip Stack Horizontal'; begin flip(false) end;

{procedure CheckForSelection;

var

x1,y1,x2,y2,LineWidth:integer;

begin

GetRoi(RoiLeft,RoiTop,RoiWidth,RoiHeight);

GetLine(x1,y1,x2,y2,LineWidth);

if (RoiWidth=0) or (x1>=0) then begin

PutMessage('Please make a rectangular selection.');

exit;

end;

end;}

macro '              Clear Outside Stack';

var

i:integer;

RoiLeft,RoiTop,RoiWidth,RoiHeight:integer;

begin

CheckForStack;

CheckForSelection;

for i:= 1 to nSlices do begin

SelectSlice(i);

Copy;

SelectAll;

Clear;

RestoreRoi;

Paste;

RestoreRoi;

end;

end;

procedure CropAndScale(fast:boolean; angle:real);

var

i,OldStack,NewStack:integer;

RoiLeft,RoiTop,RoiWidth,RoiHeight:integer;

N,NewWidth:integer;

ScaleFactor:real;

OneToOne:boolean;

begin

CheckForStack;

CheckForSelection;

GetRoi(RoiLeft,RoiTop,RoiWidth,RoiHeight);

SaveState;

OldStack:=PicNumber;

N:=nSlices;

ScaleFactor:=GetNumber('Scale factor(0.05..25):',1.0);

OneToOne:=ScaleFactor=1.0;

NewWidth:=round(RoiWidth*ScaleFactor);

if odd(NewWidth) then begin

NewWidth:=NewWidth-1;

ScaleFactor:=NewWidth/RoiWidth;

end;

SetNewSize(RoiWidth*ScaleFactor,RoiHeight*ScaleFactor);

MakeNewStack('NewStack');

NewStack:=PicNumber;

if not OneToOne then begin

if fast

then SetScaling('Nearest; Create New Window')

else SetScaling('Bilinear; Create New Window');

end;

SelectPic(OldStack);

for i:= 1 to N do begin

SelectSlice(i);

if OneToOne and (angle=0.0) then Duplicate('Temp')

else ScaleAndRotate(ScaleFactor,ScaleFactor,angle);

SelectAll;

Copy;

SelectPic(NewStack);

if i<>1 then AddSlice;

Paste;

SelectPic(nPics);

Dispose; {Temp}

SelectPic(OldStack);

{ DeleteSlice;}

end;

{Dispose;} {OldStack}

RestoreState;

SelectPic(NewStack);

Animate;

end;

macro ' [E]         Crop and Scale-Fast…';

begin CropAndScale(true, 0); end;

macro '              Crop and Scale-Smooth…';

begin CropAndScale(false, 0); end;

procedure Rotate(left:boolean);

var

i,OldStack,NewStack:integer;

RoiLeft,RoiTop,RoiWidth,RoiHeight:integer;

N,NewWidth:integer;

ScaleFactor,SliceSpacing:real;

OneToOne:boolean;

begin

CheckForStack;

SelectAll;

GetRoi(RoiLeft,RoiTop,RoiWidth,RoiHeight);

OldStack:=PicNumber;

SliceSpacing:=GetSliceSpacing;

N:=nSlices;

SetNewSize(RoiHeight,RoiWidth);

MakeNewStack('Stack');

if SliceSpacing>0 then SetSliceSpacing(SliceSpacing);

NewStack:=PicNumber;

SelectPic(OldStack);

for i:= 1 to N do begin

SelectSlice(1);

if left

then RotateLeft(true)

else RotateRight(true);

SelectAll;

Copy;

SelectPic(NewStack);

if i<>1 then AddSlice;

Paste;

ChoosePic(nPics);

Dispose;

SelectPic(OldStack);

DeleteSlice;

end;

Dispose;

end;

macro '              Rotate Left';  begin rotate(true) end;

macro '              Rotate Right'; begin rotate(false) end;

macro '              Rotate…';

var

angle:real;

begin

angle:=GetNumber('Angle(-180.0°..180.0°):',45.0);

CropAndScale(false, angle);

end;

procedure DoReslicing(horizontal:boolean);

var

OutputSpacing,stack1,stack2,width,height:integer;

RoiLeft,RoiTop,RoiWidth,RoiHeight,loc,max:integer;

InputSpacing:real;

FirstTime:boolean;

begin

CheckForStack;

CheckForSelection;

SaveState;

SetBackground(0);

SetBackground(255);

stack1:=PicNumber;

InputSpacing:=GetSliceSpacing;

if InputSpacing<=0 then InputSpacing:=1;

InputSpacing:=GetNumber('Input Slice Spacing:',InputSpacing);

SetSliceSpacing(InputSpacing);

OutputSpacing:=round(InputSpacing+0.25);

OutputSpacing:=round(GetNumber('Output Slice Spacing:',OutputSpacing));

FirstTime:=true;

GetRoi(RoiLeft,RoiTop,RoiWidth,RoiHeight);

if horizontal then begin

loc:=RoiTop+OutputSpacing;

max:=RoiTop+RoiHeight;

end else begin

loc:=RoiLeft+OutputSpacing;

max:=RoiLeft+RoiWidth;

end;

while loc<max do begin

ChoosePic(stack1);

if horizontal

then MakeLineRoi(RoiLeft,loc,RoiLeft+RoiWidth,loc)

else MakeLineRoi(loc,RoiTop,loc,RoiTop+RoiTop+RoiHeight);

Reslice;

SelectAll;

Copy;

GetPicSize(width,height);

Dispose;

if FirstTime then begin

SetNewSize(width,height);

MakeNewStack(OutputSpacing:1:2);

SetSliceSpacing(OutputSpacing);

stack2:=PicNumber;

end;

ChoosePic(stack2);

if not FirstTime then AddSlice;

Paste;

loc:=loc+OutputSpacing;

FirstTime:=false;

end;

SelectPic(stack1);

RestoreRoi;

SelectPic(stack2);

KillRoi;

RestoreState;

end;

macro '[H]         Reslice Horizontally';

begin DoReslicing(true) end;

macro '[V]         Reslice Vertically';

begin DoReslicing(false) end;

macro '(-' begin end;

macro '              Invert LUT';

var

i:integer;

begin

for i:=1 to 254 do begin

RedLUT[i]:=255-RedLut[i];

GreenLUT[i]:=255-GreenLut[i];

BlueLUT[i]:=255-BlueLut[i];

end;

UpdateLUT;

end;

macro '              Log Tranform';

var

i,v:integer;

ln255:real;

BEGIN

RedLUT[255]:=0;

GreenLUT[255]:=0;

BlueLUT[255]:=0;

ln255:=ln(255);

for i:=1 to 255 DO begin

v:=round(ln(i)*255.0/ln255);

RedLUT[255-i]:=v;

GreenLUT[255-i]:=v;

BlueLUT[255-i]:=v;

end;

UpdateLUT;

END.

macro '[X]         Gamma Tranform…';

var

i,v:integer;

n,mode,min,max:integer

gamma,mean:real;

begin

gamma:=GetNumber('Gamma(0.1-3.0):',2);

measure;

GetResults(n,mean,mode,min,max);

ShowMessage('min=',min:1,'\max=',max:1);

for i:=1 to 254 DO begin

if (i>min) and (i<max)

then v:=exp(gamma*ln((i-min)/(max-min)))*255 {x^y=exp(y*ln(x)}

else begin

if i<=min then v:=0 else v:=255;

end;

RedLUT[i]:=255-v;

GreenLUT[i]:=255-v;

BlueLUT[i]:=255-v;

end;

UpdateLUT;

end;

macro '              Reset LUT';

begin

ResetGrayMap;

end;

macro '              Set Pixels Red…';

var

v1,v2,i:integer;

begin

v1:=GetNumber('Starting Pixel Value(1-254)',10);

v2:=GetNumber('Ending Pixel Value(1-254)',10);

if v2<v1 then begin

PutMessage('Ending value less than starting value.');

exit;

end;

for i:=v1 to v2 do begin

RedLUT[i]:=255;

GreenLUT[i]:=0;

BlueLUT[i]:=0;

end;

end;

UpdateLUT;

end;

macro '              Nearly Gray LUT…';

{

Here is a macro that changes the LUT to make the values near 128 fairly visible when making polygon and line selections which use XOR drawing mode.

Play around with it to get better results. It was written on the

(incorrect) assumption that brightness = r+g+b.

j is i xor 255 and also white is 255,255,255 not 0,0,0.

{The brightness of each pixel is not quite right, there is a better way to get different colors with same brightness...)

--Edward J. Huff (huff@mcclb0.med.nyu.edu)

}

var

i,j,d: integer;

begin

while (d < 1) or (d > 63) do

d := GetNumber('Amount of color',20);

for i := d*2 to 127 do begin

j := 255 - i;

RedLUT[i] := j + d;

GreenLUT[i] := j + d;

BlueLUT[i] := j - d*2;

RedLUT[j] := i - d*2;

GreenLUT[j] := i + d;

BlueLUT[j] := i + d;

end;

UpdateLUT;

end;

macro '[U]          Move Slice Up';

var

lower,upper:integer;

begin

GetThresholds(lower,upper);

lower:=lower-1;

upper:=upper-1;

if lower<1 then lower:=1;

if lower>254 then lower:=254;

if upper<lower then upper:=lower;

if upper>254 then upper:=254;

SetDensitySlice(lower,upper);

ShowMessage(lower:4,upper:4)

end;

macro '[Y]          Move Slice Down ';

var

lower,upper:integer;

begin

GetThresholds(lower,upper);

lower:=lower+1;

upper:=upper+1;

if lower<1 then lower:=1;

if lower>254 then lower:=254;

if upper<lower then upper:=lower;

if upper>254 then upper:=254;

SetDensitySlice(lower,upper);

ShowMessage(lower:4,upper:4)

end;